TURNER FLUOROMETER CALIBRATION

Flourometers measure the intensity of chlorophyll fluorescence emission which is proportional to the chlorophyll concentration.  Calibration provides the quanitfication of the proportion.  The approach presented here is based upon first principles and notes from John Cullen and Richard Davis (Dalhousie University). Our goal is to quantify both chlorophyll a and its degraded form, pheophytin a. Both flurescence but the fluorescence bands overlap so we have to separate the effects of one from the other. This involves quantifying the fluorescence response of both in their pure form. Once done we can measure the fluorescence of a mixed sample (Xchl +Ypheo), then acidify it which converts existing chlorophyll to pheophytin (Xpheo + Ypheo). Two equations and two unknowns is easily solved, allowing retrieval of original chl and pheo concentrations. See last section on theoretical considerations for more details.

QUIK OVERVIEW OF CALIBRATION

The fluorometer is calibrated against a known concentration of chlorophyll a. A dilution series of this standard is used to assess linearity over all sensitivies. The concentration of the standard is determined spectrophotometrically. The specific steps are

1. Making the chl a standard

2. Making the working stock solution and dilution series

3. Measuring the dilutions in the spectrophotometer and calculating the chlorophyll a concentration.

4. Measuring the dilutions in the fluorometer before and after acidification; measuring blanks on all scales.

5. Entering data into the Turner_Calibration_template, saving to appropriate file name and printing out results for the calibration notebook.

1. MAKING THE CHLOROPHYLL a STANDARD

General Information:  This step is performed only occassionally. The standard is likely to go bad before it is used up. Check that it is still pure by running in the spectrophotometer. Keep stored in amber glass bottle in the   -80C freezer to help prevent evaporation. Wrap the lid with parafilm for storage. Only remove briefly from freezer to make the working stock. 

A. Purchase the Sigma chlorophyll a from Anacystis nidulans that comes in a 1mg sealed ampule.  Make sure that you purchase the algae chlorophyll a and not the spinach chl a as the spinach may contain trace amounts of chlorophyll b which will severely impact your calibration.  

B. Start by making up a fresh batch of 90% actone (500 ml). Pour 450 ml acetone in a 500 ml graduated cylinder, bring to volume 500 ml by adding MilliQ water and stirring until mixed. If it is hot in the lab (summertime), refrigerate the acetone until ready for the next steps.

C. Empty the Sigma vial contents into a 500 ml volumetric flask (although it says 1 mg on the package, contents can vary from 0.5 to 1.5 mg).  Squirt a couple drops of 90% acetone into the vial to remove any chlorophyll residue and pour into flask.  Bring to 500 ml volume with fresh cool 90% actone. Keep lights dim and minimize the amount of time that the standard is at room temperature.

2. MAKING THE WORKING STOCK STOLUTION AND DILUTION SERIES

General:  Ensure that you have sufficient fresh 90% acetone to make the dilutions (~400 ml), run the blanks (~5ml + ~40ml) and for rinsing cuvette in between samples. Assuming 1 mg, the final concentration of the standard will be 2 mg/l or 2000 μg/l. The starting concentration for the calibration should be about 100μg/l or 2.5 ml of standard in a final volume of 50 ml of 90% acetone.

A. In a 50-ml volumetric flask, add 2.5 ml of the chlorophyll a standard stock. Bring up to volume with fresh 90% acetone. Pour into a dark glass bottle.  This is the 1:1 or most concentrated of the dilutions, D1.  Measure it in the fluorometer. If it reads mid-range (3-8) on the lowest sensitivity (x1, x1). If it is too concentrated, add 10 ml acetone to the bottle and remeasure. If it reads on a more sensitive scale, add 0.5 ml more stock to the bottle and remeasure. Continue to adjust as needed.

B. Once the 1:1 dilution has been read in the fluorometer and is of correct concentration, make the dilution series: 1:2, 1:4, … 1:256, and 1:512 (~100 to 0.075 μg/l, approximately 10 dilutions).  Using a 25 ml volumetric flask, measure 25 ml of the prior dilution (D1), pour in the 50 ml volumetric, bring to volume with 90% acetone. This is now D2. Pour 25 ml into the 25 ml volumetric for the next dilution (D3) and 25 ml into the amber glass sample bottle for spectrophotometric then fluorometric readings (in this order as fluorometer sample is acidified). Repeat until all dilutions are made.

C. Then measure in the spectrophotometer (350-800nm) in 10 cm cuvettes. Keep lids on the cuvettes to minimize evaporation

1. For the baseline scan leave the reference side empty and use 90% acetone in the sample side. 

2. Run a zero in the same configuration

3. Run a blank (new acetone in the sample cuvette) 

4. For the sample readings, leave the reference side empty and place 25 ml sample in the sample side. relative to 90% acetone (see spectrophotometer instruction sheets if necessary). 

D. Each dilution will be measured in the spectrophotometer and fluorometer. There will be dilutions for which they can only be read in the spectrophotometer (concentrated) or only in the fluorometer (dilute). You need good readings in the spec to determine the concentration, you are interested in calibrating your fluorometer over its entire range. Thus you will need to optimize standard concentration based upon your instruments. Performing good dilutions is essential for the range of samples that cannot be read in the spectrophotometer.

3. SPECTROPHOTOMETRIC DETERMINATION OF CHLOROPHYLL CONCENTRATION

We use the spectrophotometer (Varian Cary3E) to measure the optical density (OD) of our dilutions from 350 to 800nm.  We are interested in the values at the red peak(rp) (at about 662-665nm) and values around 750nm (for offset correction).  You then calculate the real chlorophyll a concentration based on the specific extinction coefficient (87.67 l g-1 cm-1 at 664.3nm in 90% acetone (Jeffrey & Humphrey, 1975)).  The specific extinction coefficient is the absorbance of 1 g per liter solution at (max in a 1 cm cuvette.  


[Chl]=ODrp/87.67  

where [Chl] is in g l-1, OD is in cm-1, and the  ext. coeff. is in l g‑1 cm‑1.  If the spectrophotometric measurements were done in a 10 cm cuvette, the OD must be corrected, OD/10 cm. Convert [Chl] to (g l-1 by multiplying by 106 μg g‑1.   This is done automatically in the TURNER_CALIBRATION_TEMPLATE.XLS

In the spectrophotometer, we use quartz sample and reference cuvettes.   They should be handled carefully:  with the 10 cm cuvettes*, only touch the inlet/outlet “valves” and with the 1 cm cuvette, only touch on the frosted sides. To clean the cuvettes, use only lens paper (do not use KimWipes as they will scratch the optical surfaces).  Use the 1 cm cuvette for high chl a concentrations.  Use the 10 cm cuvettes for low concentrations.  Rinse with Milli-Q water and ethanol when finished with the cuvettes.   (*If you use the 10cm cuvettes, remember to account for this in the formula:  [Chl]=ODrp/(87.67*10).)

4. FLUOROMETRIC READINGS

General: The fluorometer excitation wavelength is a broad band excitation generally centered at 430nm.  The emission wavelength for chlorophyll a is about 668nm in 90% acetone.  The wavelength bands vary from instrument to instrument. This will impact not only your retrieved fluorescence efficiencies but also your acid ratio, Fo/Fa will range from about 2.2 for broad band filters (~650-750nm) to about 6-8 for narrow filter (~660-680nm). 

Measurements will be made on all scales. Note the large scale sensitivity is controlled by a lever that toggles between 1 and 100 while the fine scale sensitivity is controlled by the manual step switch (min 3.16 10 31.6). Thus the full range of the scale is from 1 (min*1) to 3160 (31.6*100). Record both numbers on all readings. After calibration, verify linearity not only the finescale but also over the large scale. Some instruments require two sets of calibration coefficients, one for each of the large scales. Our Turner is pretty linear over the whole ranges

Each instrument has a blank signal inherent to its electronics and configurations. This blank can change with temperature, age of lamp and any jostling. Measuring the blank response on all scales is done for each usage. This is termed the blank reading, Rblank.

Fluorescence measurements (readings) of each dilution or sample are made before and after acidification with 10% HCl, Rb and Ra. These readings are corrected for blank signal measured on the same sensitivity and scaled according to that sensitivity, resulting in the initial and acidified fluorescence signals:

Fo=(Rb-Rblank)/scale

Fa=(Ra-Rblank)/scale

where Rblank is the blank for the associated scale.

5.ENTERING DATA INTO THE TURNER_CALIBRATION_TEMPLATE FILE

Open up the TURNER_CALIBRATION_TEMPLATE.XLS file. There are five worksheets in this file:  INSTRUCTIONS, SPECTROPHOTOMETER, FLUOROMETER, F in vivo, and CALIBRATION_COEFFICIENTS. The first worksheet is INSTRUCTIONS:

Act #   Sheet
Activity

1            file

save file as Turner_Calibration_mmddyy.xls

2         SPECTROPHOTOMETRY


Measure the optical density of the dilutions in the spectrophotometer (350-800 nm) making sure you save the files in ascii.


Open the Cary files and paste the data into the worksheet in the correct column and row (see bold in example). D1 is the most concentrated.


Enter the cuvette length for each observation in cm (1 or 10)


Evaluate the spec data in terms of retrieved chl and OD spectrum to decide which observations to use in the calibration. Spectra should all have same shape and low noise.

3         FLUORESCENCE


Enter the calibration date

Enter the date  the stock solution was made (on bottle)

Enter in cell B6 the volume of stock solution in first dilution

Enter in cell C6 the volume of 90% acetone in the first dilution

Enter in subsequent B cells the volume of the previous dilution used in subsequent dilutions

Enter in subsequent C cells the volume of 90% acetone used in subsequent dilutions

Column E is the measured chlorophyll concentrations for each dilution that was measurable in the spectrophotometer. Verify the associated spectrum for each dilution.

Column F is the concentration of the D1 solution based upon the OD and dilution for each solution. They should all be the same concentration. Variations will be caused by sloppy dilutions and/or signal to noise ratios for the spec. Evaluate the dilutions to be used in the calibration

Column G is the computed chlorophyll concentration for each dilution based upon the best estimate of the D1 concentration.

Enter fluorometery readings in columns N:S for each dilution

Scaled fluorescence values are automatically computed in columns I:L

Calibration coefficents are computed from the slope of Fo-Fa vs Chl and from the median value of Fo/Fa. If one or more values are off the linear pattern, exclude from slope calculation and median calculation.**see below

Print Fluorescence worksheet, print area is set to include values and plots, place in notebook for reference

4
CALIBRATION_COFFICIENTS


M and A and calibration date are summarized

** Look at the plot of Fo-Fa vs Chl for the dilutions series. Note how well the relationship holds at the extremes and determine if those ranges should be excluded from use. Quantify the slope of the relationship. The calibration coefficient is M = 1/slope. In this case M=137.4 (mg chl/F)



[image: image1.wmf]y = 0.0073x + 0.0002

R

2

 = 0.9999

0.0001

0.001

0.01

0.1

1

10

0.01

0.1

1

10

100

1000

chl

Fo-Fa

 
Plot Fo/Fa, the acid ratio A, versus scale to determine the range overwhich the values are reliable. In this case the final two readings on the 3160 scale are noisy and thus the computation of A should exclude them. The value computed from the median of the remaining 6 points is 7.45.

6. USING THE CALIBRATION COEFFICIENTS

The calibration coefficients are now ready to use. Follow instructions in the Chlorophyll_a_analysis.doc document. The templates for chlorophyll analysis should link directly to the calibration file with the closest date. To calculate [Chl] and [Pheo], from a sample in the fluorometer.  The equations are:

 [Chl] = M * EQ \F([-  EQ \F(Ra-Rblank,scale) ],Vfilt) 
 *Vacoh

[Pheo] = M *  EQ \F([A*-  EQ \F(Rb-Rblank,scale) ],Vfilt) 
 *Vacoh

Where Vfilt is the volume of sample filtered in (ml) and Vacoh is the volume of 90% acetone in which the sample was extracted in (ml). Note that the units are mg m-3 or μg l-1.

Some useful References

Holm-Hansen, O., C.J. Lorenzen, R.W. Holmes, and J.D. Strickland. 1965. Fluorometric determination of chlorophyll. J. Cons. Cons. Int. Explor. Mer 30: 3-15.

Jeffrey, S. W. and Humphrey, G. F. (1975). New spectrophotometric equations for determining chlorophylls a, b, c1 and c2 in higher plants, algae and natural phytoplankton. Biochemie Und Physiologie Der Pflanzen, 167:191-194. 

Jeffrey, S. W., Mantoura, R. F. C. and Wright, S. W. (eds.) (1997). Phytoplankton pigments in oceanography: guidelines to modern methods. Paris: UNESCO Publishing, 661 pp.

Lorenzen, C. J. 1967. Determination of Chlorophyll and Pheo-Pigments: Spectrophotometric Equations. Limnol. Oceanogr. 12(2): 343-346.

Lorenzen, C. J. and Jeffrey, S. W. (1980). Determination of chlorophyll in seawater. Unesco Technical Papers in Marine Science, 35: 1-21.

Welschmeyer, N.A. 1994. Fluorometric analysis of chlorophyll a in the presence of chlorophyll b and pheopigments. Limnol. Oceanogr 39: 1985-1992.

Yentsch, C.S. and D.W. Menzel. 1963. A method for the determination of phytoplankton chlorophyll and pheophytin by fluorescence. Deep-Sea Res. 10: 221-231.
 More theoretical details and the first principle approach

Consider the following mixed solutions knowing that the fluorescence efficiency for chlorophyll is 1 (F/g, Fchl) and pheophytin is 0.2 (F/g, Fpheo).

If our solution has X grams chl and Y grams pheo, then Fo is the fluorescence of Xchl +Ypheo and Fa is the fluorescence of Xpheo + Ypheo.

	Chl
	Pheo
	Chl+Pheo
	Chl/Peho
	Fo
	Fa
	Fo-Fa
	Fo/Fa

	1
	1
	2
	1
	1.2
	0.4
	0.8
	3.0

	2
	2
	4
	1
	2.4
	0.8
	1.6
	3.0

	5
	5
	10
	1
	6.0
	2.0
	4.0
	3.0

	10
	10
	20
	1
	12.0
	4.0
	8.0
	3.0

	0.1
	1
	1.1
	0.1
	0.3
	0.2
	0.1
	1.4

	0.2
	2
	2.2
	0.1
	0.6
	0.4
	0.2
	1.4

	0.5
	5
	5.5
	0.1
	1.5
	1.1
	0.4
	1.4

	1
	10
	11
	0.1
	3.0
	2.2
	0.8
	1.4

	1
	0.1
	1.1
	10
	1.0
	0.2
	0.8
	4.6

	2
	0.2
	2.2
	10
	2.0
	0.4
	1.6
	4.6

	5
	0.5
	5.5
	10
	5.1
	1.1
	4.0
	4.6

	10
	1
	11
	10
	10.2
	2.2
	8.0
	4.6


Note that there are only two meaningful relationships between fluorescence and pigment values:  total chlorophyll concentration is linearly related to Fo‑Fa (and the slope is 0.8, equal to the difference between the chl and pheo fluorescence efficiencies) and total pigment concentration is linearly related to Fa (and the slope is 0.2, the pheo fluorescence efficiency). 
[image: image3.wmf]0.0

2.0

4.0

6.0

8.0

10.0

0

5

10

15

Chl (mg/m3)

Fo-Fa



 EMBED Excel.Chart.8 \s [image: image4.wmf]0.0

1.0

2.0

3.0

4.0

0

10

20

30

Chl+Pheo (mg/m3)

Fa


Since we are trying to determine the fluorescence efficiencies our calibrations are based upon relating Fo‑Fa to chlorophyll (because we are correcting the total fluorescence for that due to pheophytin). From above we see that Fo‑Fa is the fluorescence of Xchl - Xpheo. Given the fluorescence efficiencies of 1 and 0.2 in the example, that would give a constant slope of 0.8 when plotted against chlorophyll concentration.

The second important relationship that applies is that the acidified fluorescence (Fa) is due to both original pheo and acidified chl (Xpheo + Ypheo). This can be converted to the fluorescence equivalent of chl using the ratio of the chl and pheo fluorescence efficiencies. This is called the acid ratio, A (1/0.2 = 5) so that A*Fa = Xchl + Ychl, and A*Fa - Fo is the fluorescence of all the pigment in chl form minus the fluorescence of all the pigment in its original form, Ychl - Ypheo . Given the fluorescence efficiencies of 1 and 0.2 in this example, that would give a constant slope of 0.8 when plotted against pheo concentration. 


[image: image5.wmf]0

2

4

6

8

10

0

5

10

15

Pheo (mg/m3)

A*Fa-Fo


So in order to retrieve chl and pheo concentrations from an unknown sample we need to know the slope of the relationship between chlorophyll and Fo‑Fa, and the ratio Fo/Fa for a pure solution of chlorophyll. In order to ensure that our fluorometers are linear over the entire range we perform a dilution series. Thus the resulting equations are

Chl ≈ M * (Fo- Fa) 

where M is computed from the slope of Chl:(Fo-Fa) over the dilution series (or (Fchl - Fpheo)-1) with units of mg chl/F

and 

Pheo ≈  A*Fa - Fo

where A is the ratio of Fchl:Fpheo and is dimensionless.

Dilution Chart

	Dilution ID
	Approx

Concentration
	ml of DN-1
	(
	ml 90% acetone
	(

	D1
	100
	2.5 stock
	
	47.5 ml
	

	D2
	50
	25 ml D1
	
	25 ml
	

	D3
	25
	25 ml D2
	
	25 ml
	

	D4
	12.5
	25 ml D3
	
	25 ml
	

	D5
	6.25
	25 ml D4
	
	25 ml
	

	D6
	3.125
	25 ml D5
	
	25 ml
	

	D7
	1.56
	25 ml D6
	
	25 ml
	

	D8
	0.75
	25 ml D7
	
	25 ml
	

	D9
	0.375
	25 ml D8
	
	25 ml
	

	D10
	0.185
	25 ml D9
	
	25 ml
	


PAGE  
11:11
updated 6/12/06 C.Roesler

_1211720675.xls
Chart1

		127.5850662791

		63.7925331395

		12.7585066279

		6.379253314

		1.2758506628

		0.6379253314

		0.1275850663

		0.0637925331



chl

Fo-Fa

0.9256329114

0.4737974684

0.092721519

0.0431962025

0.0092594937

0.0046721519

0.0009177215

0.0004746835



cal

		March 30 2003										10 cm cells

		serial dilution of Sigma Chl (stored in CCMP freezer)																																										137.3658635236

		ID		Vol sln		Vol Acoh		dilution		OD663		OD725		chl ug/l				ODest chl		Fest chl		Fest pheo		Fo/Fa		Fo-Fa		Scale		Ro		Blank		Scale		Ra		Blank		Fo		Fa		F/chl		F/pheo

										10 cm						127.59						0.31		7.45		137.260499495																		0.0083711833		0.0010913543		7.6704545455

		1		3		57		1		0.1131693348		0.0012391931		127.67		127.67		127.59		12.71		0.61		7.16		0.93		3.16		3.4		0		31.6		4.8		0.05		1.0759		0.1503		0.0084331921		0.0011781665		7.1578947368

		2		25		25		0.5		0.0559293479		4.06E-05		63.75		127.50		63.79		6.50		-0.19		7.65		0.47		10		5.45		0		31.6		2.3		0.05		0.5450		0.0712		0.0085433196		0.0011161578		7.6542222222

		3		10		40		0.1										12.76		1.27		-0.10		7.98		0.09		31.6		3.4		0.05		316		4.6		0.4		0.1060		0.0133		0.0083091745		0.0010417473		7.9761904762

		4		5		5		0.05										6.38		0.59		0.01		7.35		0.04		100		5.1		0.1		316		2.55		0.4		0.0500		0.0068		0.0078379079		0.0010665508		7.3488372093

		5		4		16		0.01										1.28		0.13		0.01		7.17		0.01		316		3.8		0.4		1000		2.8		1.3		0.0108		0.0015		0.0084331921		0.0011756862		7.1729957806

		6		10		10		0.005										0.64		0.06		0.00		7.42		0.00		1000		6.7		1.3		3160		6.5		4.2		0.0054		0.0007		0.0084649405		0.0011409613		7.4191304348

		7		4		16		0.001										0.13		0.01		-0.00		10.67		0.00		3160		7.4		4.2		3160		4.5		4.2		0.0010		0.0001		0.0079371219		0.0007441052		10.6666666667

		8		10		10		0.0005										0.06		0.01		-0.00		16.00		0.00		3160		5.8		4.2		3160		4.3		4.2		0.0005		0.0000		0.0079371219		0.0004960701		16

																						7.20		0.0063042766





cal

		



&A

Page &P

chl

Fo-Fa



OD

		



&A

Page &P

scale

Fo-Fa



Finvivo

		538.2000684385

		269.1000342192

		134.5500171096

		67.2750085548

		33.6375042774

		20.6997544089

		16.0210427959

		8.010521398

		4.005260699

		2.0026303495

		1.0013151747

		0.5006575874



&A

Page &P

meas Chl

est Chl

544.5089383462

294.6692737414

153.3827995341

77.4836662936

38.8052144689

19.8713121115

15.4967332587

7.7800572904

3.6051296023

1.7777827042

0.9022331204

0.5002053942



Sheet4

		zero_acoh				Chl1		Chl2		Chl3

		Wavelength (nm)		Abs		Abs		Abs		Abs

		350		0.0063626086		0.0331389494		0.0144853266		-0.0032947653		0.2928262279		0.2589933041		-0.3440597129

		351		0.0066823554		0.0337790437		0.0152225485		-0.0026904682		0.2984823032		0.2721746112		-0.2809552886

		352		0.0063110213		0.0343132541		0.0149697559		-0.0028602274		0.3032027547		0.267654755		-0.2986825864

		353		0.0065864171		0.035688065		0.0159707181		-0.0029013196		0.3153510185		0.2855516599		-0.3029736958

		354		0.0064050336		0.0359276235		0.0158831775		-0.0027854387		0.3174678333		0.2839864601		-0.2908726895

		355		0.0063150665		0.0367846787		0.016406266		-0.0024634018		0.3250410437		0.293339126		-0.2572436086

		356		0.0064474363		0.0381049477		0.0170311704		-0.0020552382		0.3367073574		0.3045122298		-0.2146206551

		357		0.0065289699		0.0392575525		0.0176320933		-0.0021388077		0.3468921377		0.3152565505		-0.2233474913

		358		0.0064026695		0.040107809		0.0179326497		-0.0022425002		0.3544052731		0.3206304093		-0.2341757043

		359		0.0067532319		0.0414219722		0.0186210591		-0.001773085		0.3660176345		0.3329389639		-0.1851564718

		360		0.0069384696		0.0422664955		0.0186691042		-0.0017431269		0.3734801093		0.3337979953		-0.1820280635

		361		0.0067374045		0.0430863909		0.0194344483		-0.0019624273		0.3807249642		0.3474821193		-0.2049287631

		362		0.0063128602		0.0440889671		0.0197203811		-0.0018561933		0.3895840443		0.3525945122		-0.1938351548

		363		0.0066746534		0.0456231758		0.0203016289		-0.0012237678		0.4031407968		0.3629870476		-0.1277933787

		364		0.0066737072		0.0462876819		0.0207041018		-0.0014643826		0.4090125827		0.3701831428		-0.1529198679

		365		0.0066943429		0.0475868285		0.0215585623		-0.0012944899		0.4204922522		0.3854606412		-0.1351786263

		366		0.0072340039		0.0494628772		0.0225698836		-0.0006426353		0.4370696116		0.4035427637		-0.0671079401

		367		0.0068470505		0.0507353768		0.0229446366		-0.0008769073		0.4483138199		0.4102432357		-0.0915720656

		368		0.006724759		0.0515073948		0.0231299494		-0.0009607556		0.4551356147		0.4135565715		-0.1003280235

		369		0.0069450191		0.0537316389		0.0253611654		9.83E-05		0.4747897391		0.4534500466		0.0102651075

		370		0.0067010201		0.0549729765		0.0260424409		0.0001312617		0.4857585895		0.4656310488		0.013707156

		371		0.007078818		0.0564393178		0.0265987515		0.0007806454		0.4987156447		0.4755777164		0.0815198094

		372		0.0072587198		0.0580190159		0.0277752671		0.0011032363		0.5126743569		0.4966134632		0.1152067445

		373		0.0072689592		0.0592453331		0.0282243136		0.0011183665		0.5235104826		0.50464228		0.1167867281

		374		0.0069234511		0.0601345152		0.0283672642		0.0011281507		0.5313675765		0.5071981934		0.117808459

		375		0.0071638124		0.0614098832		0.0290920529		0.0012886496		0.5426371314		0.5201571994		0.1345687436

		376		0.0073622051		0.0620532595		0.0295383912		0.0014517272		0.5483222073		0.528137593		0.1515983137

		377		0.0067888061		0.0619977787		0.0294227265		0.0011092832		0.5478319619		0.5260695431		0.1158381933

		378		0.0070389584		0.062775895		0.0297508519		0.0012875333		0.5547076433		0.5319363276		0.1344521721

		379		0.0070695304		0.0632137805		0.030065991		0.0013461353		0.5585769381		0.5375709198		0.1405717501

		380		0.0071684439		0.063312836		0.0301994495		0.0014952336		0.5594522235		0.5399571184		0.1561415188

		381		0.0071280757		0.0637229383		0.0304936152		0.0016821137		0.5630760171		0.5452167129		0.1756566937

		382		0.0069729015		0.0634137467		0.0300800297		0.0014257034		0.560343902		0.5378219287		0.1488807447

		383		0.0068736644		0.0631551295		0.029819468		0.0013722843		0.5580586796		0.5331631634		0.1433023876

		384		0.0072780144		0.0633984134		0.0303414967		0.001633366		0.5602084122		0.5424968795		0.1705661527

		385		0.0071341284		0.06338083		0.0303584021		0.0016921186		0.56005304		0.5427991424		0.1767014617

		386		0.007377265		0.0632439926		0.0304459948		0.0020730996		0.5588439019		0.5443652746		0.2164858496

		387		0.007266643		0.062912941		0.0300722327		0.0017773125		0.5559186249		0.5376825201		0.185597934

		388		0.0069699553		0.06231457		0.029481981		0.0008557266		0.5506312294		0.5271289953		0.0893602503

		389		0.0071860766		0.0620795675		0.0293082837		0.0012292766		0.5485546731		0.5240233404		0.128368651

		390		0.0070838174		0.0621869117		0.0297017321		0.0014750513		0.5495032008		0.5310580801		0.154033954

		391		0.0069564614		0.061982166		0.0298029706		0.0014048491		0.5476940032		0.532868194		0.146703011

		392		0.0071095773		0.0618509911		0.0297805984		0.0015482523		0.5465349002		0.5324681852		0.1616780564

		393		0.0070247515		0.0621226951		0.0298663545		0.0014157566		0.548935763		0.53400148		0.1478420429

		394		0.0071298652		0.0627287254		0.0303116292		0.0014904282		0.5542908375		0.5419628573		0.1556397101

		395		0.0068829479		0.0631949753		0.0306862071		0.0016940937		0.5584107689		0.5486601988		0.1769077141

		396		0.0069075911		0.0640979633		0.0308829434		0.001652932		0.5663898563		0.5521777849		0.1726093592

		397		0.0070196479		0.0655749589		0.0317723081		0.0020450358		0.5794410565		0.5680793583		0.213555257

		398		0.0070487452		0.0671511069		0.0326540656		0.0024403837		0.593368398		0.5838449201		0.2548399163

		399		0.0071524438		0.0693407655		0.0338404626		0.0028820559		0.6127169131		0.6050573421		0.3009620604

		400		0.0071725757		0.0719268694		0.0351819731		0.0027702542		0.6355685449		0.6290431494		0.2892870286

		401		0.0067546782		0.0740060285		0.035907913		0.002892297		0.653940651		0.642022737		0.3020314947

		402		0.0070010489		0.0770745426		0.037602637		0.0033422289		0.6810550117		0.6723238942		0.3490161643

		403		0.007079476		0.0802737251		0.0391416773		0.0037463319		0.7093239987		0.6998414752		0.3912150883

		404		0.0070278826		0.0828142762		0.0403400771		0.0039305002		0.7317731112		0.7212685051		0.4104470826

		405		0.0069124834		0.0856430754		0.0418211706		0.0040787379		0.7567692747		0.7477500131		0.4259269759

		406		0.0070482455		0.0880795717		0.0430644937		0.004435719		0.7782989259		0.7699802578		0.463205141

		407		0.007000654		0.0901654512		0.0441853926		0.0048211571		0.7967304158		0.7900215941		0.5034549643

		408		0.0067726881		0.0918241218		0.0449572988		0.0048144823		0.8113869533		0.8038230464		0.5027579416

		409		0.0069482545		0.0932991132		0.0456621088		0.0048604212		0.8244204437		0.8164248386		0.5075551632

		410		0.0069893426		0.0943133682		0.0460982099		0.0050634211		0.8333827213		0.824222196		0.5287536667

		411		0.0069949455		0.0948503017		0.0465275459		0.0051626386		0.8381272357		0.8318985948		0.5391145721

		412		0.0071093403		0.0955780521		0.0467913523		0.0052649179		0.8445578678		0.8366153748		0.5497952016

		413		0.0069454927		0.0954174027		0.0467215851		0.0052294158		0.8431383188		0.8353679574		0.546087862

		414		0.0072105266		0.0953380465		0.0466528423		0.005409182		0.8424371029		0.8341388571		0.5648601588

		415		0.0067665619		0.0938457474		0.0458790101		0.0052586296		0.8292506762		0.820302969		0.5491385395

		416		0.0071643386		0.0933086127		0.0462292284		0.0053023356		0.8245043842		0.8265647668		0.5537025935

		417		0.0073758699		0.0935599208		0.0457069427		0.0055166841		0.8267250219		0.8172264548		0.5760861911

		418		0.0068793711		0.0924363136		0.0452758893		0.004907595		0.816796474		0.8095193501		0.5124813442

		419		0.0067896475		0.092536211		0.0453946739		0.0050778254		0.8176791988		0.8116431832		0.5302578527

		420		0.0068943095		0.0931252912		0.0459887832		0.0052822637		0.8228844972		0.8222656797		0.5516065657

		421		0.0070326966		0.0944798142		0.0464900583		0.0053936658		0.8348534905		0.8312283277		0.5632398511

		422		0.00701094		0.0963877514		0.047556445		0.0056804135		0.8517126267		0.8502950027		0.593183815

		423		0.0070712143		0.0989458188		0.0489395		0.006053518		0.8743165188		0.8750236122		0.6321456976

		424		0.0070973942		0.101841934		0.0501192287		0.0061670351		0.8999075075		0.8961168075		0.6439998479

		425		0.00717626		0.1057505161		0.0523534678		0.0067870445		0.9344449739		0.9360643344		0.7087450519

		426		0.0070877904		0.1093767658		0.0542823262		0.0071756546		0.966487662		0.9705517464		0.749326124

		427		0.0070246202		0.1133921742		0.0562550314		0.0074330084		1.0019690793		1.0058231237		0.7762005916

		428		0.00723074		0.117572464		0.0583253987		0.0080478145		1.0389074409		1.0428406719		0.840402435

		429		0.0071469699		0.1206941605		0.0597585328		0.0085730823		1.0664917375		1.0684646803		0.8952541385

		430		0.0072196596		0.1236300543		0.0614932626		0.0086252028		1.0924342227		1.099481131		0.9006968864

		431		0.0071124453		0.1253319681		0.0620385371		0.0089344122		1.1074728708		1.1092304751		0.9329864332

		432		0.0069485963		0.1253379434		0.0620234273		0.0086324112		1.1075256705		1.1089603169		0.9014496359

		433		0.0071322336		0.1241731197		0.061576806		0.0088870367		1.0972329202		1.1009748612		0.9280391967

		434		0.0071783918		0.1212722734		0.0599949248		0.0084856208		1.0716001257		1.0726912988		0.8861208759

		435		0.0070760818		0.1165553257		0.0577188842		0.0079552652		1.0299196855		1.031996374		0.8307378714

		436		0.0073061027		0.1100220606		0.0543306992		0.0074250819		0.9721896907		0.9714166394		0.7753728589

		437		0.0071101035		0.1022810936		0.0505176634		0.006607811		0.9037880604		0.9032407007		0.6900283847

		438		0.0069205314		0.0930121392		0.0459653027		0.0058149775		0.8218846507		0.8218458554		0.6072358185

		439		0.0072960462		0.0834339634		0.0411435626		0.0051142317		0.7372488626		0.7356345855		0.5340596263

		440		0.0071857609		0.0731842667		0.036229074		0.0039151674		0.6466793042		0.6477649996		0.408845934

		441		0.006989737		0.0635925606		0.0313099548		0.0030518591		0.5619239585		0.559812621		0.3186939535

		442		0.0072748293		0.0542573966		0.0271561053		0.0023183634		0.4794355		0.4855430343		0.2420978086

		443		0.0070522181		0.0457263552		0.0224920157		0.0014584282		0.4040525223		0.4021505093		0.1522980716

		444		0.0073972233		0.0378988124		0.0189895537		0.0007894885		0.3348858815		0.3395275361		0.0824432586

		445		0.007051534		0.0309996642		0.0153506091		6.61E-05		0.2739228279		0.274464296		0.0069044284

		446		0.0070892903		0.0253759306		0.0128531484		-0.0003709945		0.2242297406		0.2298104462		-0.0387415345

		447		0.0071311281		0.0203398764		0.0102354167		-0.0011134801		0.1797295747		0.1830061871		-0.1162764583

		448		0.0071483646		0.0162930526		0.0084055671		-0.0014028739		0.1439705607		0.1502890248		-0.1464967464

		449		0.0072281607		0.0132118948		0.0068978863		-0.0015737559		0.1167444768		0.1233321422		-0.1643412966

		450		0.0068754		0.0102977036		0.0053039077		-0.0019083084		0.0909937624		0.0948322817		-0.1992773314

		451		0.0069497013		0.0081343697		0.0043160911		-0.0020531255		0.0718778612		0.0771704165		-0.2144000335

		452		0.0070714774		0.0065777446		0.0040526604		-0.0018667611		0.0581230298		0.0724603555		-0.1949387131

		453		0.0072846483		0.0054678712		0.0031530689		-0.0019855674		0.0483158376		0.056375929		-0.2073451957

		454		0.0071790498		0.0044341497		0.0027278701		-0.002180123		0.0391815479		0.0487735015		-0.2276618931

		455		0.0070928163		0.0034194242		0.0021660686		-0.0023886992		0.0302151124		0.0387286591		-0.2494427077

		456		0.0070518763		0.0028439343		0.0018138069		-0.0022584666		0.0251299025		0.0324303253		-0.2358430103

		457		0.0073525426		0.0023778595		0.0018010699		-0.0021561689		0.0210115183		0.0322025915		-0.2251604601

		458		0.0072014462		0.002052058		0.0015492655		-0.0022530071		0.0181326329		0.0277004028		-0.2352729048

		459		0.0070701879		0.0017998742		0.0014181719		-0.0023745398		0.0159042575		0.0253564894		-0.2479640997

		460		0.007039879		0.0014818046		0.0012811983		-0.0024465164		0.0130936941		0.022907443		-0.2554803368

		461		0.0072447956		0.0011079595		0.0011576348		-0.0023228412		0.0097902806		0.0206981633		-0.242565408

		462		0.0071706017		0.0011823965		0.0012288613		-0.0022759775		0.0104480295		0.0219716718		-0.2376716128

		463		0.0070822388		0.0008601878		0.0010237013		-0.0024898099		0.0076008913		0.0183034724		-0.2600013071

		464		0.0073059183		0.0008515248		0.0010849666		-0.0022567154		0.0075243424		0.0193988778		-0.2356601476

		465		0.0072318721		0.0008396978		0.0010152427		-0.0024648432		0.0074198354		0.0181522348		-0.2573941341

		466		0.0071305232		0.000697021		0.0010218332		-0.002323871		0.006159098		0.018270071		-0.2426729471

		467		0.0071157082		0.0005045516		0.0009154896		-0.0024182524		0.0044583776		0.0163686805		-0.2525288345

		468		0.0070989733		0.0007090516		0.0010465094		-0.0023324706		0.0062654036		0.0187112742		-0.2435709704

		469		0.0071602594		0.000817523		0.0009696829		-0.0022125232		0.007223892		0.0173376396		-0.2310453268

		470		0.0071256021		0.0006222265		0.0009489543		-0.0023010063		0.00549819		0.0169670181		-0.2402852774

		471		0.0072764088		0.0006461801		0.0011351063		-0.002279531		0.0057098512		0.0202953612		-0.2380426871

		472		0.0071547329		0.0005641108		0.0011687436		-0.0023327796		0.0049846614		0.0208967856		-0.2436032345

		473		0.0070751607		0.0004983319		0.0009525862		-0.002278398		0.0044034177		0.0170319569		-0.2379243771

		474		0.0072319247		0.000581581		0.0011071292		-0.0021343257		0.0051390332		0.0197951379		-0.2228794542

		475		0.0071856817		0.0008054636		0.0011207801		-0.0021261342		0.0071173307		0.0200392133		-0.2220240528

		476		0.0071145506		0.0005286802		0.0009957313		-0.0021257736		0.0046715849		0.0178033782		-0.221986391

		477		0.0070083621		0.0004366569		0.000799006		-0.0024751904		0.0038584381		0.0142859886		-0.2584746554

		478		0.0072153956		0.0007150151		0.0010744308		-0.0022484234		0.0063180994		0.0192105005		-0.2347942427

		479		0.0070241992		0.0007531837		0.0011950895		-0.0020712113		0.0066553694		0.0213678433		-0.2162886661

		480		0.0068443683		0.0007591997		0.0011362742		-0.002267634		0.0067085287		0.0203162425		-0.2368003344

		481		0.007088948		0.0007354733		0.0011092313		-0.0022894191		0.0064988744		0.019832723		-0.2390752619

		482		0.007093579		0.0010129334		0.0012946211		-0.0019064012		0.0089505999		0.0231474387		-0.1990781786

		483		0.0070432206		0.0010002458		0.0012484869		-0.0022573336		0.0088384879		0.0223225725		-0.2357247003

		484		0.0070732664		0.0011019647		0.0014226908		-0.0021479262		0.0097373086		0.0254372859		-0.2242997096

		485		0.006913404		0.0009797493		0.0012037074		-0.002235444		0.008657374		0.0215219284		-0.2334388561

		486		0.0069320258		0.0009895823		0.0011786328		-0.002187077		0.0087442616		0.0210736024		-0.2283880678

		487		0.0070231208		0.0010654521		0.001039685		-0.002580398		0.0094146716		0.0185892563		-0.2694610771

		488		0.0066966298		0.0011976074		0.0009519896		-0.0024866182		0.0105824372		0.0170212883		-0.2596680161

		489		0.0068685361		0.0013174694		0.0011488361		-0.0023391135		0.0116415759		0.0205408457		-0.2442646621

		490		0.0068562021		0.0015166897		0.0012452678		-0.0024997445		0.0134019491		0.0222650157		-0.2610387446

		491		0.0070241732		0.001580077		0.0016713035		-0.0020824438		0.0139620595		0.0298824069		-0.2174616278

		492		0.0065723308		0.0013866706		0.001305474		-0.0021341711		0.0122530598		0.0233414838		-0.2228633101

		493		0.0065321755		0.001693262		0.0015675027		-0.0020455508		0.0149621982		0.0280264798		-0.2136090386

		494		0.006686877		0.0017028773		0.0016172305		-0.0021066081		0.0150471623		0.0289155968		-0.219985016

		495		0.006327989		0.0014857528		0.0015733222		-0.0021191791		0.0131285808		0.0281305305		-0.2212977566

		496		0.0063599558		0.0017202115		0.0015038574		-0.0019936068		0.0152003321		0.0268885208		-0.2081847204

		497		0.0066238176		0.0016785277		0.0018040591		-0.0017928324		0.014832001		0.0322560374		-0.1872186186

		498		0.0068143904		0.0018453392		0.0021333946		-0.0015712283		0.0163060002		0.0381444568		-0.164077348

		499		0.0067692962		0.0017044627		0.0020110444		-0.0017014092		0.0150611709		0.0359568723		-0.1776716478

		500		0.0071256286		0.0020498473		0.0022213028		-0.0015096314		0.0181130982		0.0397162297		-0.1576450239

		501		0.0069961557		0.0019061485		0.0020069357		-0.0015030276		0.0168433305		0.0358834088		-0.1569554169

		502		0.0071245497		0.0019599455		0.0021884425		-0.0012599562		0.0173186981		0.039128697		-0.1315723949

		503		0.0068884711		0.0020699003		0.0020757266		-0.0013339241		0.0182902929		0.0371133709		-0.1392965863

		504		0.0069667725		0.0020105764		0.0019324871		-0.0013815605		0.0177660886		0.0345522914		-0.1442710663

		505		0.0068237521		0.0019531067		0.0017997182		-0.0014829574		0.0172582681		0.0321784235		-0.1548595592

		506		0.0068230154		0.0018937208		0.0016732265		-0.0018002567		0.0167335155		0.0299167886		-0.1879939067

		507		0.0067385351		0.0019191747		0.0018197077		-0.001762618		0.0169584339		0.0325358288		-0.1840634412

		508		0.0066319918		0.0019441099		0.0016721091		-0.0017346703		0.0171787695		0.0298968106		-0.1811449689

		509		0.0069110105		0.0021169283		0.0019232827		-0.0017400846		0.0187058477		0.0343877199		-0.1817103698

		510		0.0067201057		0.0017940777		0.0017479943		-0.0017015124		0.0158530373		0.0312536155		-0.1776824187

		511		0.0066745221		0.0021174485		0.001853866		-0.0016427191		0.0187104439		0.0331465699		-0.1715428669

		512		0.0067631439		0.0021576136		0.0019702429		-0.0014732057		0.0190653558		0.0352273541		-0.1538412322

		513		0.0069570662		0.0022462294		0.0019885758		-0.0015481692		0.0198483931		0.03555514		-0.1616693765

		514		0.0065745381		0.0020672993		0.0019529248		-0.0015856202		0.0182673101		0.0349177103		-0.1655802454

		515		0.0066250791		0.0024024828		0.0018674365		-0.001745963		0.021229097		0.0333892049		-0.1823242276

		516		0.0067924345		0.0024071161		0.0022187531		-0.0015866519		0.0212700386		0.0396706413		-0.1656879789

		517		0.0065992167		0.0026583767		0.0021001506		-0.0017178595		0.0234902565		0.0375500636		-0.1793894896

		518		0.0068949931		0.0028487807		0.0022668899		-0.0012942319		0.0251727265		0.040531313		-0.1351516746

		519		0.0068016378		0.0028734046		0.0024825051		-0.0014188711		0.0253903114		0.0443864473		-0.148167274

		520		0.0066106231		0.0030662753		0.0022977004		-0.0014749084		0.0270945771		0.0410821951		-0.1540190377

		521		0.0066502336		0.0032033958		0.0023032175		-0.0016084195		0.028306217		0.0411808402		-0.1679610829

		522		0.0066421377		0.0034462723		0.0025365558		-0.0014374478		0.0304523513		0.0453528576		-0.150107172

		523		0.0065820022		0.0036158524		0.0026881993		-0.0012616598		0.0319508147		0.0480641986		-0.1317502975

		524		0.0067594368		0.0037394126		0.0026561369		-0.0012630019		0.0330426316		0.0474909323		-0.1318904534

		525		0.0066507855		0.0039243358		0.0029007401		-0.0012913413		0.0346766713		0.0518643654		-0.1348498212

		526		0.0066321497		0.0039232387		0.0026987484		-0.0013107492		0.034666977		0.0482528131		-0.1368765187

		527		0.006399333		0.0040219082		0.0026400154		-0.0013389823		0.0355388513		0.047202686		-0.1398247873

		528		0.0065650772		0.0040292502		0.0028089425		-0.0011697117		0.0356037282		0.0502230505		-0.1221485118

		529		0.0064888713		0.0042096698		0.0029103302		-0.0010517666		0.0371979722		0.0520358328		-0.1098319531

		530		0.0063934224		0.0043205097		0.0027884124		-0.0013242724		0.0381773891		0.0498559789		-0.1382886898

		531		0.0063913208		0.0043169279		0.0027284173		-0.0011547379		0.0381457386		0.0487832844		-0.1205848547

		532		0.0063524432		0.004420002		0.002763116		-0.0012799335		0.0390565342		0.0494036877		-0.1336585514

		533		0.0061668251		0.0043653501		0.0028138405		-0.0012258327		0.038573613		0.0503106263		-0.1280090161

		534		0.0060198153		0.0042399652		0.0024772983		-0.0012980516		0.0374656725		0.0442933514		-0.1355505516

		535		0.0062376438		0.0043188366		0.0029781435		-0.001049959		0.0381626049		0.0532483147		-0.1096431943

		536		0.0060296319		0.0042224256		0.0028343978		-0.0014106146		0.037310687		0.0506781849		-0.1473050768

		537		0.0060608154		0.0044467552		0.0027774705		-0.0011979538		0.0392929336		0.0496603414		-0.1250977286

		538		0.0064643831		0.0046199705		0.0027871879		-0.0013959587		0.0408235187		0.0498340859		-0.1457746201

		539		0.0062465197		0.0046162815		0.0027679875		-0.0016837205		0.0407909218		0.0494907889		-0.1758244819

		540		0.0062750401		0.0046416339		0.0028006313		-0.0014289336		0.0410149441		0.0500744501		-0.1492180596

		541		0.0063723279		0.0043640691		0.0027731718		-0.0013211239		0.0385622933		0.049583481		-0.137959909

		542		0.0063972315		0.0045401291		0.0029093658		-0.0014410599		0.0401180152		0.0520185899		-0.1504843612

		543		0.0065267622		0.0047288476		0.0030231064		-0.001326337		0.041785592		0.0540522371		-0.1385042908

		544		0.006621636		0.0046277409		0.0031394321		-0.0012467408		0.0408921812		0.0561321055		-0.1301923664

		545		0.0065295738		0.004584129		0.0030801443		-0.0012714161		0.040506812		0.055072058		-0.1327691108

		546		0.0067503401		0.0048529082		0.0033488027		-0.0010743872		0.0428818302		0.0598755896		-0.112194142

		547		0.0066690543		0.0046317177		0.0033152048		-0.0011092972		0.040927321		0.0592748687		-0.1158396521

		548		0.0064558438		0.0046054507		0.0031194601		-0.0011864404		0.0406952173		0.0557750125		-0.1238954205

		549		0.0066660312		0.004762189		0.0035029221		-0.0011025324		0.0420802067		0.0626311993		-0.1151332323

		550		0.0066539138		0.004652597		0.0034643023		-0.0009366308		0.0411118172		0.0619406879		-0.0978087685

		551		0.0065748538		0.0047849845		0.0034096926		-0.0010014598		0.0422816351		0.0609642823		-0.1045786081

		552		0.0065837102		0.0050012516		0.003563703		-0.0008916839		0.0441926392		0.0637179437		-0.0931151351

		553		0.006489607		0.0051220111		0.0034459853		-0.0009223206		0.0452597081		0.0616131851		-0.0963144114

		554		0.0064041666		0.0050824876		0.0033624722		-0.0009029469		0.0449104663		0.060119996		-0.0942912879

		555		0.0061520431		0.0051316242		0.0034600231		-0.0012671317		0.0453446526		0.0618641772		-0.1323217039

		556		0.0064222151		0.0052110241		0.0034612233		-0.0009617887		0.0460462553		0.0618856371		-0.100435909

		557		0.0062488043		0.0054390891		0.0033499766		-0.0010689646		0.0480615103		0.0598965792		-0.1116278781

		558		0.0062399283		0.0057413685		0.0034212766		-0.0011359425		0.0507325463		0.0611714012		-0.1186221262

		559		0.0061940798		0.0061186738		0.0036605943		-0.0011056308		0.0540665353		0.0654503313		-0.1154567856

		560		0.0061704484		0.0062865173		0.0037739836		-0.0010980396		0.0555496528		0.0674776966		-0.1146640647

		561		0.0060493182		0.0065231887		0.0039385981		-0.0011241692		0.0576409562		0.0704209551		-0.1173926841

		562		0.0061451904		0.0069994177		0.0040238155		-0.001084406		0.0618490661		0.0719446169		-0.1132403689

		563		0.0057864608		0.0071755494		0.0044449242		-0.0008808856		0.0634054217		0.079473914		-0.0919875123

		564		0.0061539859		0.0077314614		0.0045368331		-0.0008667287		0.0683176361		0.0811172188		-0.0905091535

		565		0.005997532		0.0080164615		0.0047593098		-0.0009751161		0.0708359868		0.085095035		-0.1018276384

		566		0.0060569569		0.0084756697		0.0048144301		-0.0007915958		0.0748936951		0.0860805699		-0.0826633211

		567		0.006177275		0.0089301048		0.0051906947		-0.0006285222		0.0789092275		0.0928080677		-0.0656341705

		568		0.0058336048		0.0090569044		0.0053759487		-0.0007559885		0.0800296685		0.0961203543		-0.0789449796

		569		0.0058061364		0.0100325812		0.0053023878		-0.0007691155		0.0886510572		0.094805106		-0.0803157871

		570		0.0056898552		0.0101661421		0.0056946808		-0.0008533979		0.0898312441		0.1018191891		-0.0891170777

		571		0.0057484247		0.0103612458		0.0056425696		-0.0007588827		0.091555242		0.100887456		-0.07924721

		572		0.0059117163		0.0108735273		0.0058272029		-0.0006724625		0.0960819226		0.1041886442		-0.070222684

		573		0.0057182079		0.0113429688		0.0062479377		-0.0006077912		0.1002300562		0.1117112557		-0.0634693108

		574		0.0057008183		0.0113076838		0.0060650678		-0.0005764602		0.0999182666		0.1084415976		-0.0601975252

		575		0.0058653252		0.0116846468		0.0062050293		-0.0006136849		0.1032492313		0.110944067		-0.0640847673

		576		0.0058822227		0.0119754728		0.0063754539		-0.0006444447		0.1058190614		0.1139912075		-0.0672968872

		577		0.0057519921		0.0121139716		0.0062858341		-0.0006419633		0.107042881		0.1123888328		-0.0670377649

		578		0.0059043425		0.0122988215		0.0064939423		-0.0007442049		0.1086762725		0.1161097455		-0.0777144616

		579		0.0058376449		0.0125236586		0.006677282		-0.000635036		0.1106630042		0.1193878044		-0.0663143803

		580		0.0060197893		0.0128283789		0.0066223717		-0.0005645683		0.1133556097		0.1184060244		-0.0589557014

		581		0.0060512871		0.0129344529		0.0067774998		-0.000430167		0.1142929122		0.1211796679		-0.0449206931

		582		0.0059544384		0.0129595222		0.0067801815		-0.0006246449		0.1145144329		0.1212276167		-0.0652292759

		583		0.0060747806		0.0131311249		0.006619297		-0.0005377844		0.1160307686		0.1183510486		-0.0561587668

		584		0.0058824853		0.0129805384		0.0064847721		-0.0005511765		0.1147001389		0.1159457838		-0.0575572554

		585		0.0059282482		0.0128837861		0.0066569103		-0.0005408351		0.1138452049		0.1190235636		-0.056477342

		586		0.005922344		0.0128368577		0.0065073171		-0.0006092905		0.1134305307		0.1163488815		-0.0636258722

		587		0.0059157573		0.0125934435		0.0062200758		-0.0006737548		0.111279646		0.111213093		-0.0703576306

		588		0.0060184244		0.0125704473		0.0063423039		-0.0005067586		0.1110764441		0.1133984961		-0.0529188564

		589		0.0056754826		0.0121107772		0.0059729666		-0.0007705109		0.1070146539		0.1067948542		-0.0804614986

		590		0.0055845389		0.0119466018		0.0058480087		-0.0007321625		0.1055639481		0.1045606448		-0.0764569251

		591		0.0058773947		0.0119600929		0.0061975196		-0.000521341		0.1056831601		0.1108097955		-0.0544416423

		592		0.0056746174		0.0117715839		0.0058348901		-0.0006005532		0.104017435		0.1043260879		-0.0627134735

		593		0.0056793904		0.011812998		0.0060518906		-0.0005219098		0.1043833827		0.1082059917		-0.0545010404

		594		0.0058547254		0.0119768297		0.0057496573		-0.0004263914		0.1058310518		0.1028021513		-0.0445264175

		595		0.00567029		0.0120078046		0.0059008789		-0.0003486993		0.1061047554		0.105505949		-0.0364133298

		596		0.0056547644		0.0121917035		0.0060811071		-0.0005812168		0.1077297446		0.1087283744		-0.0606942463

		597		0.0056610848		0.0126121771		0.0063042189		-0.0003120726		0.1114451815		0.112717547		-0.0325885494

		598		0.0054750275		0.0129445605		0.0061425911		-0.0004811606		0.1143822266		0.1098276908		-0.0502457571

		599		0.0054219468		0.0136858746		0.00659806		-0.0004003749		0.1209327126		0.1179713383		-0.0418096145

		600		0.0053899176		0.0142123401		0.0068754791		-0.0004859183		0.1255847279		0.1229315089		-0.0507425876

		601		0.0054444103		0.0146625666		0.0071769971		-0.0002705794		0.1295630713		0.1283225608		-0.028255569

		602		0.0054828133		0.0154321948		0.0072237127		-0.0003783915		0.1363637494		0.1291578218		-0.0395139749

		603		0.0055087404		0.0161275845		0.0077548618		-0.0001901167		0.1425084324		0.1386546087		-0.0198531584

		604		0.0054706237		0.017012354		0.0083096428		0.0001300188		0.1503265352		0.1485739259		0.0135773638

		605		0.0055588693		0.0176915675		0.0085958941		5.37E-05		0.1563282807		0.1536920136		0.0056040228

		606		0.0055543333		0.0182017218		0.0088660847		9.10E-05		0.1608361652		0.1585229415		0.0095053504

		607		0.0055473591		0.0190117024		0.0093152169		0.0001349904		0.1679934094		0.1665532903		0.0140965324

		608		0.0055509508		0.0198028199		0.0097113652		0.0002464526		0.174983974		0.1736363034		0.0257360983

		609		0.0054692342		0.0202619471		0.0097216889		0.0001160623		0.1790409664		0.1738208882		0.0121199413

		610		0.005637587		0.0210983623		0.0103522846		0.0003749389		0.1864317955		0.1850957511		0.0391534283

		611		0.0053571574		0.0213187765		0.0102322362		0.0004090096		0.1883794454		0.1829493214		0.0427113091

		612		0.0052323765		0.02175273		0.0104153054		6.65E-05		0.1922139955		0.1862225428		0.006944982

		613		0.0054254327		0.022242859		0.0107047521		0.0002989046		0.1965449302		0.1913977629		0.0312134597

		614		0.0054379883		0.0223962553		0.0110956896		0.0003890073		0.1979003881		0.1983876092		0.0406225419

		615		0.0053262864		0.0226779822		0.0108439866		0.0004300493		0.2003898165		0.1938872356		0.0449084056

		616		0.0055094482		0.0229166914		0.0108586643		0.0004391703		0.2024991259		0.1941496675		0.0458608762

		617		0.0055114147		0.0229393411		0.010971427		0.0004789992		0.2026992664		0.1961658302		0.0500200478

		618		0.0052438783		0.0229420699		0.0109788598		0.0004056672		0.2027233787		0.1962987281		0.042362272

		619		0.0051972703		0.0227171481		0.0110963536		0.0004398959		0.2007358982		0.1983994819		0.0459366404

		620		0.0054638078		0.0227173939		0.0108214812		0.0005245593		0.2007380707		0.1934848457		0.0547777172

		621		0.0054097064		0.0227276497		0.0109900376		0.0004503645		0.2008286936		0.1964985826		0.0470298443

		622		0.0054363371		0.0222447384		0.0107454537		0.0004616108		0.1965615372		0.1921254951		0.0482042525

		623		0.0054697846		0.0216616206		0.0102695273		0.0002099342		0.1914089239		0.1836160748		0.0219226206

		624		0.0053207045		0.0213188306		0.0101611074		0.0003635653		0.1883799228		0.1816775592		0.0379657266

		625		0.0052827089		0.0209740829		0.009910184		0.0003107687		0.1853336233		0.177191123		0.0324523811

		626		0.0052697915		0.020278953		0.0096662864		0.0001674109		0.1791912364		0.1728303078		0.0174820828

		627		0.0053277276		0.0197816864		0.0093487799		0.0002164865		0.1747972311		0.1671533868		0.0226068589

		628		0.0052110241		0.0191721581		0.0091549233		0.0001999634		0.1694112466		0.1636872878		0.0208814084

		629		0.00525842		0.0183744635		0.0088097611		-3.73E-06		0.1623625656		0.1575158925		-0.0003892552

		630		0.0052906484		0.0180197414		0.0083946148		-4.31E-05		0.1592281286		0.1500932		-0.004497856

		631		0.0055256765		0.0175960772		0.0086155646		9.06E-05		0.1554844976		0.1540437156		0.0094620904

		632		0.0053404374		0.0167816021		0.0080025652		0.0001270669		0.1482875381		0.1430834713		0.0132691113

		633		0.005270394		0.0163939819		0.0078695361		-0.0003290413		0.1448624038		0.1407049493		-0.0343605225

		634		0.0051209894		0.016154021		0.0075034015		-0.0001936874		0.1427420339		0.1341585728		-0.020226032

		635		0.0053580222		0.0159642994		0.0075504156		-0.0003171943		0.141065594		0.1349991707		-0.0331233881

		636		0.0053817406		0.0159255508		0.0074302694		-5.63E-05		0.1407231991		0.1328509927		-0.0058817219

		637		0.0053559258		0.0158863738		0.0075308718		-0.0001414697		0.1403770186		0.1346497334		-0.0147731346

		638		0.005306161		0.0160572566		0.007682241		-6.36E-05		0.141886993		0.1373561704		-0.0066385058

		639		0.005404071		0.016342802		0.0078550912		5.76E-05		0.144410162		0.1404466804		0.0060176612

		640		0.0053970469		0.016942367		0.0079654418		6.53E-05		0.1497081077		0.1424197152		0.0068206171

		641		0.0052469959		0.0174250193		0.0083763255		-9.10E-05		0.1539729762		0.1497661925		-0.0094979538

		642		0.0054860641		0.0187848043		0.0088258497		0.0002911076		0.1659884659		0.1578035518		0.0303992515

		643		0.0056288545		0.0199712124		0.0095744235		0.0005596267		0.1764719427		0.1711878273		0.0584396754

		644		0.0051608048		0.0210323054		0.0100498265		0.0004522562		0.1858480963		0.1796878893		0.0472273802

		645		0.005377023		0.0228840001		0.0111296298		0.0004928377		0.2022102554		0.1989944497		0.0514651577

		646		0.0052400269		0.0251631848		0.0120711448		0.0006934171		0.2223498521		0.2158284555		0.0724108908

		647		0.0054124584		0.0277853124		0.0135149611		0.001307603		0.2455197994		0.2416434597		0.1365479689

		648		0.0052145873		0.0306354538		0.0148159061		0.0013567037		0.2707045492		0.2649039659		0.1416753693

		649		0.0052227611		0.0341744945		0.0168385375		0.0017795478		0.301976631		0.3010680112		0.1858313568

		650		0.0053499243		0.0385811031		0.0188106466		0.0022179724		0.3409148172		0.3363287313		0.2316143625

		651		0.0053474871		0.0434504412		0.021343464		0.0026610855		0.3839418274		0.3816147487		0.2778869594

		652		0.0052836263		0.0486990549		0.0241278913		0.0030371046		0.4303202361		0.4313994747		0.3171531999

		653		0.0054140049		0.0549574383		0.0271717571		0.0038848687		0.4856212892		0.4858228838		0.405681955

		654		0.0052475203		0.0614894144		0.0303864963		0.0041053393		0.5433398944		0.5433014597		0.4287048547

		655		0.005402341		0.0687817931		0.0340711661		0.0053690821		0.6077776567		0.6091822517		0.560672674

		656		0.0055206954		0.076306358		0.037750259		0.0060023875		0.6742670894		0.6749633328		0.6268063312

		657		0.0052399482		0.083548598		0.0415460356		0.0066377744		0.7382618104		0.7428306809		0.6931573521

		658		0.0053845448		0.0910115466		0.0450553521		0.0073214765		0.8042067819		0.8055762112		0.7645537392

		659		0.0054763905		0.0975896567		0.0485645123		0.0081332615		0.8623330417		0.868318944		0.8493253362

		660		0.0052579218		0.1036967337		0.0510311574		0.0085986927		0.9162971037		0.9124218199		0.8979285386

		661		0.0055363472		0.1081333011		0.0532441176		0.0090779178		0.955500015		0.9519888874		0.9479721766

		662		0.0055034449		0.111479871		0.0549660735		0.0091981776		0.985071364		0.9827769422		0.9605304491

		663		0.0052795648		0.1131693348		0.0559293479		0.0095761437		1		1		1

		664		0.0052842028		0.1129908338		0.0555895828		0.0094461506		0.9984227088		0.9939251017		0.9864253213

		665		0.0056410488		0.111578092		0.0548411459		0.0093885874		0.9859392758		0.9805432739		0.9804142184

		666		0.0054255114		0.108061716		0.0531676039		0.0090956548		0.9548674664		0.9506208446		0.9498243876

		667		0.0052308831		0.1034891307		0.0507930145		0.0084453411		0.914462658		0.9081638961		0.8819146211

		668		0.0052096359		0.0977738053		0.0479481109		0.0081179384		0.8639602365		0.8572978709		0.8477252087

		669		0.0053832605		0.0907426029		0.0445824042		0.0072503495		0.8018303112		0.7971200434		0.7571262225

		670		0.00557232		0.0827998668		0.040414393		0.0064513511		0.731645785		0.7225972497		0.6736898804

		671		0.0054786974		0.0744119808		0.036311809		0.0060609207		0.6575277741		0.6492442761		0.6329187247

		672		0.0052274507		0.0665585846		0.0323717929		0.0048694522		0.5881326835		0.5787979687		0.5084982396

		673		0.0056028678		0.0585775711		0.0287717264		0.0045924219		0.5176099264		0.5144298567		0.4795690285

		674		0.0058890185		0.0510437042		0.0249876399		0.0038961261		0.451038298		0.4467715218		0.4068575181

		675		0.0054918313		0.0441859625		0.0214391835		0.0032914428		0.3904411262		0.3833261841		0.3437127576

		676		0.0054794839		0.0376591794		0.0182584133		0.0024987247		0.3327684084		0.326454965		0.260932251

		677		0.0054059317		0.032113947		0.0153196091		0.0017403013		0.2837689824		0.2739100251		0.1817329936

		678		0.0054739001		0.0271605961		0.0131955631		0.0012757982		0.239999609		0.2359327189		0.1332267173

		679		0.0054267435		0.022996299		0.0110744992		0.0009707207		0.2032025638		0.1980087316		0.10136864

		680		0.0058452799		0.019697275		0.0094596958		0.0008337325		0.1740513457		0.1691365292		0.0870634953

		681		0.0057302471		0.0165592562		0.0078128669		0.000458579		0.1463228203		0.1396917225		0.0478876468

		682		0.0057044639		0.0136848865		0.0065068179		0.0002559841		0.1209239812		0.1163399562		0.0267314398

		683		0.0055692522		0.0118703237		0.0056337584		0.0003236175		0.1048899307		0.1007299136		0.0337941338

		684		0.0058448599		0.0099907005		0.0048333546		-9.73E-05		0.0882809865		0.0864189339		-0.0101573845

		685		0.005913212		0.008843259		0.0040129991		-0.0002274781		0.0781418303		0.0717512232		-0.0237546635

		686		0.0055891275		0.0076358183		0.0031817518		-0.0003583728		0.0674725034		0.0568887702		-0.0374234936

		687		0.0056866291		0.0064300178		0.0028176184		-0.0005092924		0.0568176686		0.0503781742		-0.0531834554

		688		0.0058201198		0.0056945495		0.0026926273		-0.0004601122		0.0503188389		0.0481433694		-0.048047758

		689		0.0059322631		0.0053633689		0.0025610058		-0.0005242885		0.0473924222		0.0457900165		-0.0547494344

		690		0.0057926523		0.0048543741		0.001971907		-0.0008198626		0.0428947834		0.0352571066		-0.0856151091

		691		0.005901299		0.0042982087		0.0019222167		-0.0005559335		0.0379803306		0.0343686598		-0.0580540068

		692		0.0059026107		0.0042067948		0.0016774102		-0.0002918439		0.0371725679		0.0299915925		-0.0304761372

		693		0.0058515244		0.0041881329		0.0018605212		-0.000361373		0.0370076659		0.0332655613		-0.0377368018

		694		0.0056446414		0.0036588192		0.001627		-0.0008881707		0.0323304829		0.0290902739		-0.0927482609

		695		0.005611626		0.0031419352		0.0010933229		-0.0008768039		0.0277631323		0.0195482858		-0.0915612765

		696		0.0058149775		0.0031494705		0.0013304779		-0.0006545251		0.0278297169		0.0237885463		-0.0683495451

		697		0.0059100632		0.0036630481		0.0013135488		-0.0003961852		0.0323678508		0.0234858588		-0.0413721002

		698		0.00583085		0.0030450292		0.0011256334		-0.00095404		0.0269068405		0.0201259877		-0.0996267459

		699		0.0059312135		0.0029621928		0.0012714625		-0.0006043274		0.0261748711		0.0227333696		-0.0631075911

		700		0.0059436001		0.0032079071		0.0010578749		-0.0005411454		0.0283460806		0.0189144862		-0.0565097399

		701		0.0061487611		0.0032542502		0.0012540425		-0.000164293		0.0287555826		0.0224219045		-0.0171564932

		702		0.0061611538		0.0028265032		0.0011245173		-0.0003450782		0.0249758758		0.0201060326		-0.0360351894

		703		0.0059697386		0.0029498651		0.0008476084		-0.0005830264		0.0260659398		0.0151549841		-0.0608832177

		704		0.0058907242		0.0029297192		0.0012810426		-0.0003815468		0.0258879242		0.022904658		-0.039843473

		705		0.0057736333		0.0029189561		0.000905762		-0.0008405835		0.0257928184		0.0161947538		-0.0877789172

		706		0.0059722317		0.0032688815		0.0009285379		-0.0005477639		0.0288848695		0.0166019801		-0.0572008847

		707		0.006048216		0.0029759542		0.0010514915		-0.0005180833		0.0262964717		0.0188003529		-0.0541014523

		708		0.0063055321		0.0030867662		0.0008608881		-0.0003480785		0.0272756417		0.0153924216		-0.0363485067

		709		0.0055913562		0.002779268		0.0009108463		-0.000595228		0.0245584899		0.0162856589		-0.0621573847

		710		0.0061876993		0.0029817405		0.0011457475		-0.0004707658		0.0263476013		0.0204856222		-0.0491602729

		711		0.0063944729		0.0030362704		0.0014850515		-0.0002070384		0.0268294445		0.0265522758		-0.0216202262

		712		0.0062225703		0.0026437659		0.0008384789		-0.0004199779		0.0233611507		0.0149917509		-0.0438566839

		713		0.0061043138		0.0028867985		0.0007955569		-0.0002129375		0.0255086634		0.0142243188		-0.022236251

		714		0.0061894585		0.0034748702		0.0007809047		-0.0008616134		0.0307050513		0.0139623427		-0.0899749896

		715		0.0063420413		0.002531192		0.0010081334		-0.0005438342		0.0223664126		0.0180251235		-0.0567905196

		716		0.006474236		0.0030768856		0.0008773067		-0.0003816503		0.0271883333		0.0156859819		-0.0398542773

		717		0.0060514705		0.0026810886		0.0007942083		-0.0003198328		0.0236909463		0.014200207		-0.03339891

		718		0.0065214806		0.0030334292		0.0010330942		-0.0004354425		0.0268043384		0.0184714159		-0.045471591

		719		0.0065735658		0.002763585		0.001207056		5.41E-05		0.0244199099		0.0215817998		0.0056526861

		720		0.0065387976		0.0030639551		0.001069163		3.50E-05		0.0270740755		0.0191163139		0.0036521332

		721		0.0063608224		0.0028101669		0.0005377255		-0.0003723394		0.0248315228		0.0096143702		-0.038881979

		722		0.0066931597		0.0025632191		0.0011065581		0.0002176779		0.0226494134		0.0197849283		0.0227312671

		723		0.0066128573		0.0028570406		0.0012858976		6.21E-05		0.0252457136		0.0229914636		0.0064826708

		724		0.0067295176		0.0027259425		0.0011874582		-0.0001089145		0.0240872895		0.0212313966		-0.0113735233

		725		0.006789411		0.0029421765		0.0012391931		4.06E-05		0.0259980015		0.0221564024		0.0042387725

		726		0.0064367959		0.0029843468		0.0010828128		0.0002738045		0.0263706315		0.0193603685		0.0285923581

		727		0.005698143		0.0021576397		0.0003557671		-0.0002273228		0.0190655862		0.0063610095		-0.0237384539

		728		0.007211132		0.0028728314		0.0011473048		0.0001389781		0.0253852462		0.020513466		0.0145129516

		729		0.006281422		0.0024101096		0.0004331847		-0.0006103245		0.0212964901		0.0077452123		-0.0637338489

		730		0.0064352457		0.0024468654		0.0004323814		-0.0002359124		0.0216212761		0.0077308501		-0.0246354286

		731		0.0067762905		0.0029944079		0.0007776372		-9.37E-05		0.0264595343		0.013903921		-0.0097844289

		732		0.0072501651		0.0029630528		0.000952716		5.80E-05		0.026182471		0.0170342767		0.0060555106

		733		0.0063014873		0.0019616617		0.0005521621		-0.0004978641		0.0173338629		0.0098724924		-0.05199004

		734		0.0070779235		0.0020814748		0.0005949822		-0.0007085419		0.0183925686		0.0106381033		-0.0739903153

		735		0.0070485347		0.0025460075		0.0005887093		6.55E-05		0.0224973268		0.0105259453		0.0068422456

		736		0.006880397		0.0025289007		0.0007369253		-7.78E-05		0.022346166		0.0131760046		-0.008119603

		737		0.0067300699		0.0023044925		8.48E-05		-0.0003339559		0.020363224		0.0015154643		-0.0348737311

		738		0.0065715946		0.0019305891		0.0001588395		-0.0011212259		0.0170592952		0.0028400021		-0.1170853235

		739		0.0069476757		0.0024287996		0.0009035831		-0.0003761672		0.021461641		0.0161557949		-0.0392817039

		740		0.0076891966		0.0029523671		0.0008813272		-0.0001140386		0.0260880483		0.0157578665		-0.0119086142

		741		0.0066509433		0.0019740653		0.000689165		-0.0003852711		0.0174434649		0.0123220644		-0.0402323874

		742		0.0062568928		0.0018876371		0.000176138		-0.0004078231		0.0166797576		0.0031492939		-0.0425874039

		743		0.0068921791		0.0020832175		0.0001509673		-0.0006529742		0.018407968		0.0026992504		-0.0681875982

		744		0.0065091564		0.002220236		0.0006102761		-0.0008149535		0.0196187068		0.0109115545		-0.0851024748

		745		0.0071538645		0.0021520725		0.0004594082		-0.0005503492		0.0190163926		0.0082140813		-0.0574708631

		746		0.006710405		0.0018008359		0.0006959061		-0.0003958748		0.0159127554		0.012442593		-0.0413396932

		747		0.0067334874		0.0018421937		0.0002357557		-0.0005179282		0.0162782052		0.0042152418		-0.0540852533

		748		0.0071191029		0.0016704979		-6.76E-05		-0.0008623885		0.0147610476		-0.0012088267		-0.0900559235

		749		0.0066871662		0.0018098298		0.0001283875		-0.0009795584		0.0159922284		0.0022955299		-0.1022915361

		750		0.0069605908		0.001956461		0.0006032512		-0.0007054923		0.0172879075		0.0107859512		-0.0736718616

		751		0.0070273825		0.0015882349		0.0001328153		-0.0008145918		0.0140341455		0.0023746984		-0.0850647036

		752		0.0074894694		0.001999602		0.0003402749		8.50E-05		0.0176691147		0.0060840131		0.0088807915

		753		0.0064644879		0.0015002729		8.65E-05		-0.0011115693		0.0132568853		0.0015460174		-0.1160769287

		754		0.0069904211		0.0017833686		-0.0005354576		-0.0008849673		0.0157584088		-0.0095738211		-0.0924137482

		755		0.0070873168		0.001239297		-0.0004034266		-0.0003882712		0.0109508197		-0.0072131466		-0.0405456774

		756		0.0067464751		0.0016687829		-0.0001592213		-0.0006932421		0.014745893		-0.0028468291		-0.0723926192

		757		0.0063916096		0.0009305095		-0.0004732482		-0.001465363		0.0082222757		-0.008461536		-0.1530222403

		758		0.0072735134		0.0011379872		0.0004465294		0.0001687575		0.0100556149		0.0079838129		0.0176227021

		759		0.0074483869		0.0011821369		0.000260206		-0.0004542164		0.0104457355		0.0046524059		-0.0474320826

		760		0.007149891		0.0008606028		-0.0007375376		-0.0007403286		0.007604558		-0.0131869521		-0.0773096764

		761		0.0072166063		0.0010392958		-0.0004869009		-0.0003884781		0.0091835461		-0.0087056421		-0.0405672831

		762		0.0072568511		0.0009642607		-0.0010191763		-0.00122511		0.0085205119		-0.0182225677		-0.1279335467

		763		0.0064382409		0.0009065921		0.0001232865		-0.0007752137		0.0080109339		0.0022043253		-0.0809525971

		764		0.0073088668		0.0018207735		5.99E-05		-5.70E-05		0.0160889297		0.0010706074		-0.005957401

		765		0.0072203963		0.0012238252		-0.0003937024		-0.0005653955		0.0108141057		-0.0070392808		-0.0590420877

		766		0.0071475487		0.0013026439		-0.0007416724		-0.0008927689		0.0115105732		-0.0132608797		-0.0932284365

		767		0.0073841107		0.0012023835		8.92E-05		-0.0004895383		0.0106246409		0.0015946246		-0.0511206095

		768		0.0069052242		0.0010458088		-0.0009646815		-0.0011969212		0.0092410968		-0.0172482165		-0.1249898977

		769		0.0072223176		0.0010267891		-0.0005688597		-0.0004608363		0.0090730327		-0.010171042		-0.0481233672

		770		0.0062258793		0.000951341		-0.0004638359		-0.0006722558		0.0084063497		-0.0082932465		-0.0702010935

		771		0.005964437		0.0002266649		-0.0011822582		-0.0007689088		0.0020028828		-0.0211384236		-0.0802941967

		772		0.0068462356		0.001437936		-0.0001560643		-0.0004506996		0.0127060566		-0.0027903837		-0.0470648346

		773		0.0070445095		4.37E-05		-0.0007037866		-0.0007150545		0.0003861268		-0.0125834941		-0.0746704035

		774		0.0072638262		0.0001245294		-0.0007171736		-0.0005460576		0.0011003807		-0.0128228497		-0.0570227024

		775		0.0064320932		4.31E-05		-0.0002446053		-0.001081824		0.0003806366		-0.004373469		-0.1129707309

		776		0.0071122614		0.0013595861		-0.0001126929		-0.0006199405		0.0120137327		-0.002014915		-0.0647380072

		777		0.0075494414		0.0009902827		0.0001169168		-0.0003490097		0.0087504512		0.0020904368		-0.0364457429

		778		0.0069488068		0.001475415		0.0001344725		-0.0002615766		0.0130372328		0.0024043286		-0.0273154406

		779		0.006690531		0.0006660646		-0.000643204		-0.0006915364		0.0058855574		-0.0115002944		-0.0722144976

		780		0.0063049281		0.0003770115		0.0004255667		0.0001857974		0.0033313929		0.0076090045		0.0194021151

		781		0.0074644508		0.0010660749		-2.81E-05		-8.15E-05		0.009420175		-0.0005026208		-0.0085087887

		782		0.0083187725		0.0011221815		0.0011102434		4.06E-05		0.0099159506		0.0198508193		0.0042360691

		783		0.0067111147		7.44E-05		-0.000170762		-0.0005849394		0.0006572165		-0.0030531742		-0.0610829844

		784		0.0063071079		0.0001662974		-0.0003402154		-0.0008197076		0.0014694565		-0.00608295		-0.0855989223

		785		0.0067987451		0.0012925442		-0.0002139725		-0.0005889722		0.0114213288		-0.0038257637		-0.0615041084

		786		0.0076478054		0.0012534194		0.0002136636		-0.0001791457		0.0110756102		0.0038202405		-0.018707501

		787		0.0071111824		0.0011909363		0.0004890282		0.0001527282		0.0105234895		0.0087436773		0.0159488193

		788		0.0066838013		0.000776574		-2.59E-05		-0.0006625893		0.0068620534		-0.0004628196		-0.0691916592

		789		0.0073030749		0.0009160344		-0.0004168746		-0.0006050511		0.0080943693		-0.007453593		-0.06318317

		790		0.0076204068		0.0012237732		0.0001810842		0.0002598952		0.0108136469		0.0032377314		0.0271398629

		791		0.0077395779		0.0009304576		-0.000197879		-0.000793973		0.0082218174		-0.0035380177		-0.0829115631

		792		0.0068267761		-0.0003762707		-0.0014471487		-0.0009316714		-0.0033248468		-0.0258745863		-0.0972908764

		793		0.0073957224		0.0007789079		-0.0006009668		0.0006569389		0.0068826764		-0.0107451065		0.0686016104

		794		0.0076335263		0.0014902204		-0.0002941203		0.0002394853		0.0131680586		-0.0052587832		0.0250085362

		795		0.0081879413		0.0002233757		0.0007871286		0.00080606		0.0019738181		0.0140736234		0.0841737617

		796		0.0067921979		0.0009154637		-0.0002899813		3.55E-05		0.0080893262		-0.0051847796		0.0037116078

		797		0.0073139477		0.0008153964		0.0012606366		0.0012093922		0.0072051005		0.0225398052		0.1262921955

		798		0.0073607308		0.0006295369		5.75E-05		-0.000262094		0.0055627867		0.0010289468		-0.0273694715

		799		0.0070814495		0.0005516437		-0.0003098998		-0.0004577333		0.0048744977		-0.0055409151		-0.0477993306

		800		0.0075571854		0.0008047893		-0.0002264949		-0.0001808535		0.0071113726		-0.0040496615		-0.0188858368

		801		0.0066532562		0.0001499315		-0.0007645676		-0.0007855495		0.0013248423		-0.0136702395		-0.082031927

		802		0.0063765836		0.0005169915		0.0007961015		-0.0009661795		0.0045683006		0.0142340559		-0.1008944273

		803		0.0060943114		0.000118548		-0.0003110379		-0.0003266616		0.0010475278		-0.0055612651		-0.0341120161

		804		0.0060736258		0.0011510941		-0.0009035153		-0.0006177691		0.0101714311		-0.0161545825		-0.0645112585

		805		0.0061042611		0.0005787038		0.0003818305		-0.0004371492		0.0051136098		0.0068270155		-0.0456498188

		806		0.0060338578		-9.86E-05		0.0003161309		-0.0007016157		-0.0008713875		0.0056523264		-0.07326704

		807		0.0062491461		0.0013461353		-0.0007523191		0.000517147		0.0118948767		-0.0134512396		0.0540036812

		808		0.0077264812		0.00085845		0.00102928		0.0010159173		0.0075855356		0.0184032186		0.1060883498

		809		0.0072704069		0.0001201534		-0.000319212		0.0006496281		0.0010617134		-0.0057074146		0.067838169

		810		0.0074872044		0.0005189612		-0.000513119		0.0005276175		0.0045857054		-0.0091744149		0.0550970765

		811		0.0062703392		-0.0004442865		-0.0013486852		-0.0004051852		-0.0039258559		-0.0241140884		-0.0423119368

		812		0.0061637005		-3.62E-06		-0.0009105932		-4.04E-05		-0.000032023		-0.0162811338		-0.0042167531

		813		0.0072767246		0.0004457262		-0.0011562353		-0.0006166318		0.0039385771		-0.0206731419		-0.0643924926

		814		0.0062044254		-0.0012136488		-0.0006892104		-0.0010025961		-0.0107241845		-0.0123228751		-0.1046972707

		815		0.0072027361		0.000302298		0.0005203607		0.0003830482		0.0026712003		0.0093038938		0.0400002593

		816		0.0068770042		0.0004183634		-0.0002064692		-1.24E-05		0.0036967909		-0.0036916069		-0.0012975043

		817		0.0075419345		0.0005322827		0.0001205418		0.0002772236		0.004703418		0.0021552512		0.0289494035

		818		0.0068811332		-0.0003589418		-0.0003165735		-0.0008003293		-0.0031717231		-0.0056602401		-0.0835753186

		819		0.0056334175		-0.0006136332		-0.0001835445		-0.0007229108		-0.0054222568		-0.0032817207		-0.0754908055

		820		0.0060566156		0.000222262		-0.0015643674		-0.0006072225		0.0019639773		-0.0279704216		-0.0634099187

		821		0.0057069291		-0.0002771503		-0.0024162959		-0.0004844186		-0.0024489874		-0.043202648		-0.0505859805

		822		0.0058076582		-0.0014865175		-0.001075575		-0.0018737719		-0.0131353382		-0.0192309593		-0.1956708204

		823		0.0072564823		-0.0005109989		-0.0021017136		0.0001288018		-0.0045153476		-0.0375780094		0.0134502781

		824		0.00757349		0.0012203986		-0.000130601		-0.0001775932		0.0107838275		-0.0023351059		-0.0185453779

		825		0.0060007866		-0.0008392917		-0.0021367471		-0.0017406003		-0.0074162468		-0.038204399		-0.1817642244

		826		0.007528475		-0.0001942049		-0.000140124		0.0003683582		-0.0017160557		-0.0025053757		0.0384662345

		827		0.007321029		0.0005172766		-0.0010343096		-0.0007836374		0.0045708193		-0.0184931465		-0.0818322514

		828		0.0055899927		-0.000546523		-0.0014338874		-0.0008609417		-0.0048292496		-0.0256374776		-0.0899048448

		829		0.0050577382		-0.0018646992		-0.0023125419		-0.002317949		-0.0164770711		-0.0413475577		-0.2420545303

		830		0.006661668		-0.0019281012		-0.0020813618		-0.0009399371		-0.0170373112		-0.0372141266		-0.0981540279

		831		0.0057976102		-0.0013932236		-0.0023553336		-0.002474058		-0.0123109644		-0.0421126607		-0.258356394

		832		0.0063595879		-0.0012799851		-0.0016205407		-0.0009088883		-0.0113103525		-0.0289747835		-0.0949117226

		833		0.005823845		-0.0007023393		-0.0020932639		-0.0012344025		-0.0062060923		-0.037426932		-0.1289039273

		834		0.0055073509		-0.0023986879		-0.0026361314		-0.0031502589		-0.021195564		-0.0471332398		-0.3289694695

		835		0.0067146374		-0.0017122386		-0.001889494		-0.0016873305		-0.0151298817		-0.0337835875		-0.1762014644

		836		0.0072787781		0.0006048065		-0.0019951013		-0.0003897196		0.0053442616		-0.0356718148		-0.0406969203

		837		0.0055343024		-0.0013853795		-0.0039340351		-0.0021221158		-0.0122416507		-0.0703393697		-0.2216044241

		838		0.0054130089		-0.0026353081		-0.0034461063		-0.0027088362		-0.0232864148		-0.0616153498		-0.2828733895

		839		0.0068449993		-8.56E-05		-0.0006993931		0.000609991		-0.0007561287		-0.0125049393		0.0636990197

		840		0.00641943		-0.0003626145		-0.0017669491		-0.0016937252		-0.0032041767		-0.0315925206		-0.1768692378

		841		0.0059720217		-0.0006301248		-0.0038509129		-0.0039248001		-0.005567982		-0.0688531702		-0.4098518367

		842		0.0058392193		-0.0013133816		-0.0007781595		-0.0018928961		-0.0116054548		-0.0139132585		-0.1976678918

		843		0.0056040213		-0.0006867292		-0.0017749411		-0.0013269047		-0.0060681563		-0.0317354157		-0.1385635796

		844		0.0058772112		-7.92E-06		-0.0021475656		-0.0021986156		-0.0000699928		-0.0383978301		-0.2295930018

		845		0.0043324842		-0.0011519496		-0.0036161351		-0.0021319042		-0.0101789909		-0.0646554133		-0.2226265928

		846		0.0062730177		-0.0007499417		-0.0016013012		-0.0020398826		-0.0066267217		-0.0286307871		-0.2130171238

		847		0.0070434571		-0.0016316817		-0.0012736875		-0.000633175		-0.014418055		-0.0227731506		-0.0661200355

		848		0.006261305		-0.000886104		-0.0026596987		-0.0011295913		-0.007829895		-0.047554617		-0.1179588873

		849		0.0059828344		0.0010091712		-0.0011610372		-0.000310624		0.008917356		-0.0207589985		-0.0324372793

		850		0.0049338024		-0.0007747486		-0.002482706		-0.0019179472		-0.0068459236		-0.0443900399		-0.2002838784





Sheet4

		350		350		350		350

		351		351		351		351

		352		352		352		352

		353		353		353		353

		354		354		354		354

		355		355		355		355

		356		356		356		356

		357		357		357		357

		358		358		358		358

		359		359		359		359

		360		360		360		360

		361		361		361		361

		362		362		362		362

		363		363		363		363

		364		364		364		364

		365		365		365		365

		366		366		366		366

		367		367		367		367

		368		368		368		368

		369		369		369		369

		370		370		370		370

		371		371		371		371

		372		372		372		372

		373		373		373		373

		374		374		374		374

		375		375		375		375

		376		376		376		376

		377		377		377		377

		378		378		378		378

		379		379		379		379

		380		380		380		380

		381		381		381		381

		382		382		382		382

		383		383		383		383

		384		384		384		384

		385		385		385		385

		386		386		386		386

		387		387		387		387

		388		388		388		388

		389		389		389		389

		390		390		390		390

		391		391		391		391

		392		392		392		392

		393		393		393		393

		394		394		394		394

		395		395		395		395

		396		396		396		396

		397		397		397		397

		398		398		398		398

		399		399		399		399

		400		400		400		400

		401		401		401		401

		402		402		402		402

		403		403		403		403

		404		404		404		404

		405		405		405		405

		406		406		406		406

		407		407		407		407

		408		408		408		408

		409		409		409		409

		410		410		410		410

		411		411		411		411

		412		412		412		412

		413		413		413		413

		414		414		414		414

		415		415		415		415

		416		416		416		416

		417		417		417		417

		418		418		418		418

		419		419		419		419

		420		420		420		420

		421		421		421		421

		422		422		422		422

		423		423		423		423

		424		424		424		424

		425		425		425		425

		426		426		426		426

		427		427		427		427

		428		428		428		428

		429		429		429		429

		430		430		430		430

		431		431		431		431

		432		432		432		432

		433		433		433		433

		434		434		434		434

		435		435		435		435

		436		436		436		436

		437		437		437		437

		438		438		438		438

		439		439		439		439

		440		440		440		440

		441		441		441		441

		442		442		442		442

		443		443		443		443

		444		444		444		444

		445		445		445		445

		446		446		446		446

		447		447		447		447

		448		448		448		448

		449		449		449		449

		450		450		450		450

		451		451		451		451

		452		452		452		452

		453		453		453		453

		454		454		454		454

		455		455		455		455

		456		456		456		456

		457		457		457		457

		458		458		458		458

		459		459		459		459

		460		460		460		460

		461		461		461		461

		462		462		462		462

		463		463		463		463

		464		464		464		464

		465		465		465		465

		466		466		466		466

		467		467		467		467

		468		468		468		468

		469		469		469		469

		470		470		470		470

		471		471		471		471

		472		472		472		472

		473		473		473		473

		474		474		474		474

		475		475		475		475

		476		476		476		476

		477		477		477		477

		478		478		478		478

		479		479		479		479

		480		480		480		480

		481		481		481		481

		482		482		482		482

		483		483		483		483

		484		484		484		484

		485		485		485		485

		486		486		486		486

		487		487		487		487

		488		488		488		488

		489		489		489		489

		490		490		490		490

		491		491		491		491

		492		492		492		492

		493		493		493		493

		494		494		494		494

		495		495		495		495

		496		496		496		496

		497		497		497		497

		498		498		498		498

		499		499		499		499

		500		500		500		500

		501		501		501		501

		502		502		502		502

		503		503		503		503

		504		504		504		504

		505		505		505		505

		506		506		506		506

		507		507		507		507

		508		508		508		508

		509		509		509		509

		510		510		510		510

		511		511		511		511

		512		512		512		512

		513		513		513		513

		514		514		514		514

		515		515		515		515

		516		516		516		516

		517		517		517		517

		518		518		518		518

		519		519		519		519

		520		520		520		520

		521		521		521		521

		522		522		522		522

		523		523		523		523

		524		524		524		524

		525		525		525		525

		526		526		526		526

		527		527		527		527

		528		528		528		528

		529		529		529		529

		530		530		530		530

		531		531		531		531

		532		532		532		532

		533		533		533		533

		534		534		534		534

		535		535		535		535

		536		536		536		536

		537		537		537		537

		538		538		538		538

		539		539		539		539

		540		540		540		540

		541		541		541		541

		542		542		542		542

		543		543		543		543

		544		544		544		544

		545		545		545		545

		546		546		546		546

		547		547		547		547

		548		548		548		548

		549		549		549		549

		550		550		550		550

		551		551		551		551

		552		552		552		552

		553		553		553		553

		554		554		554		554

		555		555		555		555

		556		556		556		556

		557		557		557		557

		558		558		558		558

		559		559		559		559

		560		560		560		560

		561		561		561		561

		562		562		562		562

		563		563		563		563

		564		564		564		564

		565		565		565		565

		566		566		566		566

		567		567		567		567

		568		568		568		568

		569		569		569		569

		570		570		570		570

		571		571		571		571

		572		572		572		572

		573		573		573		573

		574		574		574		574

		575		575		575		575

		576		576		576		576

		577		577		577		577

		578		578		578		578

		579		579		579		579

		580		580		580		580

		581		581		581		581

		582		582		582		582

		583		583		583		583

		584		584		584		584

		585		585		585		585

		586		586		586		586

		587		587		587		587

		588		588		588		588

		589		589		589		589

		590		590		590		590

		591		591		591		591

		592		592		592		592

		593		593		593		593

		594		594		594		594

		595		595		595		595

		596		596		596		596

		597		597		597		597

		598		598		598		598

		599		599		599		599

		600		600		600		600

		601		601		601		601

		602		602		602		602

		603		603		603		603

		604		604		604		604

		605		605		605		605

		606		606		606		606

		607		607		607		607

		608		608		608		608

		609		609		609		609

		610		610		610		610

		611		611		611		611

		612		612		612		612

		613		613		613		613

		614		614		614		614

		615		615		615		615

		616		616		616		616

		617		617		617		617

		618		618		618		618

		619		619		619		619

		620		620		620		620

		621		621		621		621

		622		622		622		622

		623		623		623		623

		624		624		624		624

		625		625		625		625

		626		626		626		626

		627		627		627		627

		628		628		628		628

		629		629		629		629

		630		630		630		630

		631		631		631		631

		632		632		632		632

		633		633		633		633

		634		634		634		634

		635		635		635		635

		636		636		636		636

		637		637		637		637

		638		638		638		638

		639		639		639		639

		640		640		640		640

		641		641		641		641

		642		642		642		642

		643		643		643		643

		644		644		644		644

		645		645		645		645

		646		646		646		646

		647		647		647		647

		648		648		648		648

		649		649		649		649

		650		650		650		650

		651		651		651		651

		652		652		652		652

		653		653		653		653

		654		654		654		654

		655		655		655		655

		656		656		656		656

		657		657		657		657

		658		658		658		658

		659		659		659		659

		660		660		660		660

		661		661		661		661

		662		662		662		662

		663		663		663		663

		664		664		664		664

		665		665		665		665

		666		666		666		666

		667		667		667		667

		668		668		668		668

		669		669		669		669

		670		670		670		670

		671		671		671		671

		672		672		672		672

		673		673		673		673

		674		674		674		674

		675		675		675		675

		676		676		676		676

		677		677		677		677

		678		678		678		678

		679		679		679		679

		680		680		680		680

		681		681		681		681

		682		682		682		682

		683		683		683		683

		684		684		684		684

		685		685		685		685

		686		686		686		686

		687		687		687		687

		688		688		688		688

		689		689		689		689

		690		690		690		690

		691		691		691		691

		692		692		692		692

		693		693		693		693

		694		694		694		694

		695		695		695		695

		696		696		696		696

		697		697		697		697

		698		698		698		698

		699		699		699		699

		700		700		700		700

		701		701		701		701

		702		702		702		702

		703		703		703		703

		704		704		704		704

		705		705		705		705

		706		706		706		706

		707		707		707		707

		708		708		708		708

		709		709		709		709

		710		710		710		710

		711		711		711		711

		712		712		712		712

		713		713		713		713

		714		714		714		714

		715		715		715		715

		716		716		716		716

		717		717		717		717

		718		718		718		718

		719		719		719		719

		720		720		720		720

		721		721		721		721

		722		722		722		722

		723		723		723		723

		724		724		724		724

		725		725		725		725

		726		726		726		726

		727		727		727		727

		728		728		728		728

		729		729		729		729

		730		730		730		730

		731		731		731		731

		732		732		732		732

		733		733		733		733

		734		734		734		734

		735		735		735		735

		736		736		736		736

		737		737		737		737

		738		738		738		738

		739		739		739		739

		740		740		740		740

		741		741		741		741

		742		742		742		742

		743		743		743		743

		744		744		744		744

		745		745		745		745

		746		746		746		746

		747		747		747		747

		748		748		748		748

		749		749		749		749

		750		750		750		750

		751		751		751		751

		752		752		752		752

		753		753		753		753

		754		754		754		754

		755		755		755		755

		756		756		756		756

		757		757		757		757

		758		758		758		758

		759		759		759		759

		760		760		760		760

		761		761		761		761

		762		762		762		762

		763		763		763		763

		764		764		764		764

		765		765		765		765

		766		766		766		766

		767		767		767		767

		768		768		768		768

		769		769		769		769

		770		770		770		770

		771		771		771		771

		772		772		772		772

		773		773		773		773

		774		774		774		774

		775		775		775		775

		776		776		776		776

		777		777		777		777

		778		778		778		778

		779		779		779		779

		780		780		780		780

		781		781		781		781

		782		782		782		782

		783		783		783		783

		784		784		784		784

		785		785		785		785

		786		786		786		786

		787		787		787		787

		788		788		788		788

		789		789		789		789

		790		790		790		790

		791		791		791		791

		792		792		792		792

		793		793		793		793

		794		794		794		794

		795		795		795		795

		796		796		796		796

		797		797		797		797

		798		798		798		798

		799		799		799		799

		800		800		800		800

		801		801		801		801

		802		802		802		802

		803		803		803		803

		804		804		804		804

		805		805		805		805

		806		806		806		806

		807		807		807		807

		808		808		808		808

		809		809		809		809

		810		810		810		810

		811		811		811		811

		812		812		812		812

		813		813		813		813

		814		814		814		814

		815		815		815		815

		816		816		816		816

		817		817		817		817

		818		818		818		818

		819		819		819		819

		820		820		820		820

		821		821		821		821

		822		822		822		822

		823		823		823		823

		824		824		824		824

		825		825		825		825

		826		826		826		826

		827		827		827		827

		828		828		828		828

		829		829		829		829

		830		830		830		830

		831		831		831		831

		832		832		832		832

		833		833		833		833

		834		834		834		834

		835		835		835		835

		836		836		836		836

		837		837		837		837

		838		838		838		838

		839		839		839		839

		840		840		840		840

		841		841		841		841

		842		842		842		842

		843		843		843		843

		844		844		844		844

		845		845		845		845

		846		846		846		846

		847		847		847		847

		848		848		848		848

		849		849		849		849

		850		850		850		850



zero

Dil1

Dil_2

Dil_3

0.0063626086

0.0331389494

0.0144853266

-0.0032947653

0.0066823554

0.0337790437

0.0152225485

-0.0026904682

0.0063110213

0.0343132541

0.0149697559

-0.0028602274

0.0065864171

0.035688065

0.0159707181

-0.0029013196

0.0064050336

0.0359276235

0.0158831775

-0.0027854387

0.0063150665

0.0367846787

0.016406266

-0.0024634018

0.0064474363

0.0381049477

0.0170311704

-0.0020552382

0.0065289699

0.0392575525

0.0176320933

-0.0021388077

0.0064026695

0.040107809

0.0179326497

-0.0022425002

0.0067532319

0.0414219722

0.0186210591

-0.001773085

0.0069384696

0.0422664955

0.0186691042

-0.0017431269

0.0067374045

0.0430863909

0.0194344483

-0.0019624273

0.0063128602

0.0440889671

0.0197203811

-0.0018561933

0.0066746534

0.0456231758

0.0203016289

-0.0012237678

0.0066737072

0.0462876819

0.0207041018

-0.0014643826

0.0066943429

0.0475868285

0.0215585623

-0.0012944899

0.0072340039

0.0494628772

0.0225698836

-0.0006426353

0.0068470505

0.0507353768

0.0229446366

-0.0008769073

0.006724759

0.0515073948

0.0231299494

-0.0009607556

0.0069450191

0.0537316389

0.0253611654

0.0000983001

0.0067010201

0.0549729765

0.0260424409

0.0001312617

0.007078818

0.0564393178

0.0265987515

0.0007806454

0.0072587198

0.0580190159

0.0277752671

0.0011032363

0.0072689592

0.0592453331

0.0282243136

0.0011183665

0.0069234511

0.0601345152

0.0283672642

0.0011281507

0.0071638124

0.0614098832

0.0290920529

0.0012886496

0.0073622051

0.0620532595

0.0295383912

0.0014517272

0.0067888061

0.0619977787

0.0294227265

0.0011092832

0.0070389584

0.062775895

0.0297508519

0.0012875333

0.0070695304

0.0632137805

0.030065991

0.0013461353

0.0071684439

0.063312836

0.0301994495

0.0014952336

0.0071280757

0.0637229383

0.0304936152

0.0016821137

0.0069729015

0.0634137467

0.0300800297

0.0014257034

0.0068736644

0.0631551295

0.029819468

0.0013722843

0.0072780144

0.0633984134

0.0303414967

0.001633366

0.0071341284

0.06338083

0.0303584021

0.0016921186

0.007377265

0.0632439926

0.0304459948

0.0020730996

0.007266643

0.062912941

0.0300722327

0.0017773125

0.0069699553

0.06231457

0.029481981

0.0008557266

0.0071860766

0.0620795675

0.0293082837

0.0012292766

0.0070838174

0.0621869117

0.0297017321

0.0014750513

0.0069564614

0.061982166

0.0298029706

0.0014048491

0.0071095773

0.0618509911

0.0297805984

0.0015482523

0.0070247515

0.0621226951

0.0298663545

0.0014157566

0.0071298652

0.0627287254

0.0303116292

0.0014904282

0.0068829479

0.0631949753

0.0306862071

0.0016940937

0.0069075911

0.0640979633

0.0308829434

0.001652932

0.0070196479

0.0655749589

0.0317723081

0.0020450358

0.0070487452

0.0671511069

0.0326540656

0.0024403837

0.0071524438

0.0693407655

0.0338404626

0.0028820559

0.0071725757

0.0719268694

0.0351819731

0.0027702542

0.0067546782

0.0740060285

0.035907913

0.002892297

0.0070010489

0.0770745426

0.037602637

0.0033422289

0.007079476

0.0802737251

0.0391416773

0.0037463319

0.0070278826

0.0828142762

0.0403400771

0.0039305002

0.0069124834

0.0856430754

0.0418211706

0.0040787379

0.0070482455

0.0880795717

0.0430644937

0.004435719

0.007000654

0.0901654512

0.0441853926

0.0048211571

0.0067726881

0.0918241218

0.0449572988

0.0048144823

0.0069482545

0.0932991132

0.0456621088

0.0048604212

0.0069893426

0.0943133682

0.0460982099

0.0050634211

0.0069949455

0.0948503017

0.0465275459

0.0051626386

0.0071093403

0.0955780521

0.0467913523

0.0052649179

0.0069454927

0.0954174027

0.0467215851

0.0052294158

0.0072105266

0.0953380465

0.0466528423

0.005409182

0.0067665619

0.0938457474

0.0458790101

0.0052586296

0.0071643386

0.0933086127

0.0462292284

0.0053023356

0.0073758699

0.0935599208

0.0457069427

0.0055166841

0.0068793711

0.0924363136

0.0452758893

0.004907595

0.0067896475

0.092536211

0.0453946739

0.0050778254

0.0068943095

0.0931252912

0.0459887832

0.0052822637

0.0070326966

0.0944798142

0.0464900583

0.0053936658

0.00701094

0.0963877514

0.047556445

0.0056804135

0.0070712143

0.0989458188

0.0489395

0.006053518

0.0070973942

0.101841934

0.0501192287

0.0061670351

0.00717626

0.1057505161

0.0523534678

0.0067870445

0.0070877904

0.1093767658

0.0542823262

0.0071756546

0.0070246202

0.1133921742

0.0562550314

0.0074330084

0.00723074

0.117572464

0.0583253987

0.0080478145

0.0071469699

0.1206941605

0.0597585328

0.0085730823

0.0072196596

0.1236300543

0.0614932626

0.0086252028

0.0071124453

0.1253319681

0.0620385371

0.0089344122

0.0069485963

0.1253379434

0.0620234273

0.0086324112

0.0071322336

0.1241731197

0.061576806

0.0088870367

0.0071783918

0.1212722734

0.0599949248

0.0084856208

0.0070760818

0.1165553257

0.0577188842

0.0079552652

0.0073061027

0.1100220606

0.0543306992

0.0074250819

0.0071101035

0.1022810936

0.0505176634

0.006607811

0.0069205314

0.0930121392

0.0459653027

0.0058149775

0.0072960462

0.0834339634

0.0411435626

0.0051142317

0.0071857609

0.0731842667

0.036229074

0.0039151674

0.006989737

0.0635925606

0.0313099548

0.0030518591

0.0072748293

0.0542573966

0.0271561053

0.0023183634

0.0070522181

0.0457263552

0.0224920157

0.0014584282

0.0073972233

0.0378988124

0.0189895537

0.0007894885

0.007051534

0.0309996642

0.0153506091

0.0000661178

0.0070892903

0.0253759306

0.0128531484

-0.0003709945

0.0071311281

0.0203398764

0.0102354167

-0.0011134801

0.0071483646

0.0162930526

0.0084055671

-0.0014028739

0.0072281607

0.0132118948

0.0068978863

-0.0015737559

0.0068754

0.0102977036

0.0053039077

-0.0019083084

0.0069497013

0.0081343697

0.0043160911

-0.0020531255

0.0070714774

0.0065777446

0.0040526604

-0.0018667611

0.0072846483

0.0054678712

0.0031530689

-0.0019855674

0.0071790498

0.0044341497

0.0027278701

-0.002180123

0.0070928163

0.0034194242

0.0021660686

-0.0023886992

0.0070518763

0.0028439343

0.0018138069

-0.0022584666

0.0073525426

0.0023778595

0.0018010699

-0.0021561689

0.0072014462

0.002052058

0.0015492655

-0.0022530071

0.0070701879

0.0017998742

0.0014181719

-0.0023745398

0.007039879

0.0014818046

0.0012811983

-0.0024465164

0.0072447956

0.0011079595

0.0011576348

-0.0023228412

0.0071706017

0.0011823965

0.0012288613

-0.0022759775

0.0070822388

0.0008601878

0.0010237013

-0.0024898099

0.0073059183

0.0008515248

0.0010849666

-0.0022567154

0.0072318721

0.0008396978

0.0010152427

-0.0024648432

0.0071305232

0.000697021

0.0010218332

-0.002323871

0.0071157082

0.0005045516

0.0009154896

-0.0024182524

0.0070989733

0.0007090516

0.0010465094

-0.0023324706

0.0071602594

0.000817523

0.0009696829

-0.0022125232

0.0071256021

0.0006222265

0.0009489543

-0.0023010063

0.0072764088

0.0006461801

0.0011351063

-0.002279531

0.0071547329

0.0005641108

0.0011687436

-0.0023327796

0.0070751607

0.0004983319

0.0009525862

-0.002278398

0.0072319247

0.000581581

0.0011071292

-0.0021343257

0.0071856817

0.0008054636

0.0011207801

-0.0021261342

0.0071145506

0.0005286802

0.0009957313

-0.0021257736

0.0070083621

0.0004366569

0.000799006

-0.0024751904

0.0072153956

0.0007150151

0.0010744308

-0.0022484234

0.0070241992

0.0007531837

0.0011950895

-0.0020712113

0.0068443683

0.0007591997

0.0011362742

-0.002267634

0.007088948

0.0007354733

0.0011092313

-0.0022894191

0.007093579

0.0010129334

0.0012946211

-0.0019064012

0.0070432206

0.0010002458

0.0012484869

-0.0022573336

0.0070732664

0.0011019647

0.0014226908

-0.0021479262

0.006913404

0.0009797493

0.0012037074

-0.002235444

0.0069320258

0.0009895823

0.0011786328

-0.002187077

0.0070231208

0.0010654521

0.001039685

-0.002580398

0.0066966298

0.0011976074

0.0009519896

-0.0024866182

0.0068685361

0.0013174694

0.0011488361

-0.0023391135

0.0068562021

0.0015166897

0.0012452678

-0.0024997445

0.0070241732

0.001580077

0.0016713035

-0.0020824438

0.0065723308

0.0013866706

0.001305474

-0.0021341711

0.0065321755

0.001693262

0.0015675027

-0.0020455508

0.006686877

0.0017028773

0.0016172305

-0.0021066081

0.006327989

0.0014857528

0.0015733222

-0.0021191791

0.0063599558

0.0017202115

0.0015038574

-0.0019936068

0.0066238176

0.0016785277

0.0018040591

-0.0017928324

0.0068143904

0.0018453392

0.0021333946

-0.0015712283

0.0067692962

0.0017044627

0.0020110444

-0.0017014092

0.0071256286

0.0020498473

0.0022213028

-0.0015096314

0.0069961557

0.0019061485

0.0020069357

-0.0015030276

0.0071245497

0.0019599455

0.0021884425

-0.0012599562

0.0068884711

0.0020699003

0.0020757266

-0.0013339241

0.0069667725

0.0020105764

0.0019324871

-0.0013815605

0.0068237521

0.0019531067

0.0017997182

-0.0014829574

0.0068230154

0.0018937208

0.0016732265

-0.0018002567

0.0067385351

0.0019191747

0.0018197077

-0.001762618

0.0066319918

0.0019441099

0.0016721091

-0.0017346703

0.0069110105

0.0021169283

0.0019232827

-0.0017400846

0.0067201057

0.0017940777

0.0017479943

-0.0017015124

0.0066745221

0.0021174485

0.001853866

-0.0016427191

0.0067631439

0.0021576136

0.0019702429

-0.0014732057

0.0069570662

0.0022462294

0.0019885758

-0.0015481692

0.0065745381

0.0020672993

0.0019529248

-0.0015856202

0.0066250791

0.0024024828

0.0018674365

-0.001745963

0.0067924345

0.0024071161

0.0022187531

-0.0015866519

0.0065992167

0.0026583767

0.0021001506

-0.0017178595

0.0068949931

0.0028487807

0.0022668899

-0.0012942319

0.0068016378

0.0028734046

0.0024825051

-0.0014188711

0.0066106231

0.0030662753

0.0022977004

-0.0014749084

0.0066502336

0.0032033958

0.0023032175

-0.0016084195

0.0066421377

0.0034462723

0.0025365558

-0.0014374478

0.0065820022

0.0036158524

0.0026881993

-0.0012616598

0.0067594368

0.0037394126

0.0026561369

-0.0012630019

0.0066507855

0.0039243358

0.0029007401

-0.0012913413

0.0066321497

0.0039232387

0.0026987484

-0.0013107492

0.006399333

0.0040219082

0.0026400154

-0.0013389823

0.0065650772

0.0040292502

0.0028089425

-0.0011697117

0.0064888713

0.0042096698

0.0029103302

-0.0010517666

0.0063934224

0.0043205097

0.0027884124

-0.0013242724

0.0063913208

0.0043169279

0.0027284173

-0.0011547379

0.0063524432

0.004420002

0.002763116

-0.0012799335

0.0061668251

0.0043653501

0.0028138405

-0.0012258327

0.0060198153

0.0042399652

0.0024772983

-0.0012980516

0.0062376438

0.0043188366

0.0029781435

-0.001049959

0.0060296319

0.0042224256

0.0028343978

-0.0014106146

0.0060608154

0.0044467552

0.0027774705

-0.0011979538

0.0064643831

0.0046199705

0.0027871879

-0.0013959587

0.0062465197

0.0046162815

0.0027679875

-0.0016837205

0.0062750401

0.0046416339

0.0028006313

-0.0014289336

0.0063723279

0.0043640691

0.0027731718

-0.0013211239

0.0063972315

0.0045401291

0.0029093658

-0.0014410599

0.0065267622

0.0047288476

0.0030231064

-0.001326337

0.006621636

0.0046277409

0.0031394321

-0.0012467408

0.0065295738

0.004584129

0.0030801443

-0.0012714161

0.0067503401

0.0048529082

0.0033488027

-0.0010743872

0.0066690543

0.0046317177

0.0033152048

-0.0011092972

0.0064558438

0.0046054507

0.0031194601

-0.0011864404

0.0066660312

0.004762189

0.0035029221

-0.0011025324

0.0066539138

0.004652597

0.0034643023

-0.0009366308

0.0065748538

0.0047849845

0.0034096926

-0.0010014598

0.0065837102

0.0050012516

0.003563703

-0.0008916839

0.006489607

0.0051220111

0.0034459853

-0.0009223206

0.0064041666

0.0050824876

0.0033624722

-0.0009029469

0.0061520431

0.0051316242

0.0034600231

-0.0012671317

0.0064222151

0.0052110241

0.0034612233

-0.0009617887

0.0062488043

0.0054390891

0.0033499766

-0.0010689646

0.0062399283

0.0057413685

0.0034212766

-0.0011359425

0.0061940798

0.0061186738

0.0036605943

-0.0011056308

0.0061704484

0.0062865173

0.0037739836

-0.0010980396

0.0060493182

0.0065231887

0.0039385981

-0.0011241692

0.0061451904

0.0069994177

0.0040238155

-0.001084406

0.0057864608

0.0071755494

0.0044449242

-0.0008808856

0.0061539859

0.0077314614

0.0045368331

-0.0008667287

0.005997532

0.0080164615

0.0047593098

-0.0009751161

0.0060569569

0.0084756697

0.0048144301

-0.0007915958

0.006177275

0.0089301048

0.0051906947

-0.0006285222

0.0058336048

0.0090569044

0.0053759487

-0.0007559885

0.0058061364

0.0100325812

0.0053023878

-0.0007691155

0.0056898552

0.0101661421

0.0056946808

-0.0008533979

0.0057484247

0.0103612458

0.0056425696

-0.0007588827

0.0059117163

0.0108735273

0.0058272029

-0.0006724625

0.0057182079

0.0113429688

0.0062479377

-0.0006077912

0.0057008183

0.0113076838

0.0060650678

-0.0005764602

0.0058653252

0.0116846468

0.0062050293

-0.0006136849

0.0058822227

0.0119754728

0.0063754539

-0.0006444447

0.0057519921

0.0121139716

0.0062858341

-0.0006419633

0.0059043425

0.0122988215

0.0064939423

-0.0007442049

0.0058376449

0.0125236586

0.006677282

-0.000635036

0.0060197893

0.0128283789

0.0066223717

-0.0005645683

0.0060512871

0.0129344529

0.0067774998

-0.000430167

0.0059544384

0.0129595222

0.0067801815

-0.0006246449

0.0060747806

0.0131311249

0.006619297

-0.0005377844

0.0058824853

0.0129805384

0.0064847721

-0.0005511765

0.0059282482

0.0128837861

0.0066569103

-0.0005408351

0.005922344

0.0128368577

0.0065073171

-0.0006092905

0.0059157573

0.0125934435

0.0062200758

-0.0006737548

0.0060184244

0.0125704473

0.0063423039

-0.0005067586

0.0056754826

0.0121107772

0.0059729666

-0.0007705109

0.0055845389

0.0119466018

0.0058480087

-0.0007321625

0.0058773947

0.0119600929

0.0061975196

-0.000521341

0.0056746174

0.0117715839

0.0058348901

-0.0006005532

0.0056793904

0.011812998

0.0060518906

-0.0005219098

0.0058547254

0.0119768297

0.0057496573

-0.0004263914

0.00567029

0.0120078046

0.0059008789

-0.0003486993

0.0056547644

0.0121917035

0.0060811071

-0.0005812168

0.0056610848

0.0126121771

0.0063042189

-0.0003120726

0.0054750275

0.0129445605

0.0061425911

-0.0004811606

0.0054219468

0.0136858746

0.00659806

-0.0004003749

0.0053899176

0.0142123401

0.0068754791

-0.0004859183

0.0054444103

0.0146625666

0.0071769971

-0.0002705794

0.0054828133

0.0154321948

0.0072237127

-0.0003783915

0.0055087404

0.0161275845

0.0077548618

-0.0001901167

0.0054706237

0.017012354

0.0083096428

0.0001300188

0.0055588693

0.0176915675

0.0085958941

0.0000536649

0.0055543333

0.0182017218

0.0088660847

0.0000910246

0.0055473591

0.0190117024

0.0093152169

0.0001349904

0.0055509508

0.0198028199

0.0097113652

0.0002464526

0.0054692342

0.0202619471

0.0097216889

0.0001160623

0.005637587

0.0210983623

0.0103522846

0.0003749389

0.0053571574

0.0213187765

0.0102322362

0.0004090096

0.0052323765

0.02175273

0.0104153054

0.0000665061

0.0054254327

0.022242859

0.0107047521

0.0002989046

0.0054379883

0.0223962553

0.0110956896

0.0003890073

0.0053262864

0.0226779822

0.0108439866

0.0004300493

0.0055094482

0.0229166914

0.0108586643

0.0004391703

0.0055114147

0.0229393411

0.010971427

0.0004789992

0.0052438783

0.0229420699

0.0109788598

0.0004056672

0.0051972703

0.0227171481

0.0110963536

0.0004398959

0.0054638078

0.0227173939

0.0108214812

0.0005245593

0.0054097064

0.0227276497

0.0109900376

0.0004503645

0.0054363371

0.0222447384

0.0107454537

0.0004616108

0.0054697846

0.0216616206

0.0102695273

0.0002099342

0.0053207045

0.0213188306

0.0101611074

0.0003635653

0.0052827089

0.0209740829

0.009910184

0.0003107687

0.0052697915

0.020278953

0.0096662864

0.0001674109

0.0053277276

0.0197816864

0.0093487799

0.0002164865

0.0052110241

0.0191721581

0.0091549233

0.0001999634

0.00525842

0.0183744635

0.0088097611

-0.0000037276

0.0052906484

0.0180197414

0.0083946148

-0.0000430721

0.0055256765

0.0175960772

0.0086155646

0.0000906103

0.0053404374

0.0167816021

0.0080025652

0.0001270669

0.005270394

0.0163939819

0.0078695361

-0.0003290413

0.0051209894

0.016154021

0.0075034015

-0.0001936874

0.0053580222

0.0159642994

0.0075504156

-0.0003171943

0.0053817406

0.0159255508

0.0074302694

-0.0000563242

0.0053559258

0.0158863738

0.0075308718

-0.0001414697

0.005306161

0.0160572566

0.007682241

-0.0000635713

0.005404071

0.016342802

0.0078550912

0.000057626

0.0053970469

0.016942367

0.0079654418

0.0000653152

0.0052469959

0.0174250193

0.0083763255

-0.0000909538

0.0054860641

0.0187848043

0.0088258497

0.0002911076

0.0056288545

0.0199712124

0.0095744235

0.0005596267

0.0051608048

0.0210323054

0.0100498265

0.0004522562

0.005377023

0.0228840001

0.0111296298

0.0004928377

0.0052400269

0.0251631848

0.0120711448

0.0006934171

0.0054124584

0.0277853124

0.0135149611

0.001307603

0.0052145873

0.0306354538

0.0148159061

0.0013567037

0.0052227611

0.0341744945

0.0168385375

0.0017795478

0.0053499243

0.0385811031

0.0188106466

0.0022179724

0.0053474871

0.0434504412

0.021343464

0.0026610855

0.0052836263

0.0486990549

0.0241278913

0.0030371046

0.0054140049

0.0549574383

0.0271717571

0.0038848687

0.0052475203

0.0614894144

0.0303864963

0.0041053393

0.005402341

0.0687817931

0.0340711661

0.0053690821

0.0055206954

0.076306358

0.037750259

0.0060023875

0.0052399482

0.083548598

0.0415460356

0.0066377744

0.0053845448

0.0910115466

0.0450553521

0.0073214765

0.0054763905

0.0975896567

0.0485645123

0.0081332615

0.0052579218

0.1036967337

0.0510311574

0.0085986927

0.0055363472

0.1081333011

0.0532441176

0.0090779178

0.0055034449

0.111479871

0.0549660735

0.0091981776

0.0052795648

0.1131693348

0.0559293479

0.0095761437

0.0052842028

0.1129908338

0.0555895828

0.0094461506

0.0056410488

0.111578092

0.0548411459

0.0093885874

0.0054255114

0.108061716

0.0531676039

0.0090956548

0.0052308831

0.1034891307

0.0507930145

0.0084453411

0.0052096359

0.0977738053

0.0479481109

0.0081179384

0.0053832605

0.0907426029

0.0445824042

0.0072503495

0.00557232

0.0827998668

0.040414393

0.0064513511

0.0054786974

0.0744119808

0.036311809

0.0060609207

0.0052274507

0.0665585846

0.0323717929

0.0048694522

0.0056028678

0.0585775711

0.0287717264

0.0045924219

0.0058890185

0.0510437042

0.0249876399

0.0038961261

0.0054918313

0.0441859625

0.0214391835

0.0032914428

0.0054794839

0.0376591794

0.0182584133

0.0024987247

0.0054059317

0.032113947

0.0153196091

0.0017403013

0.0054739001

0.0271605961

0.0131955631

0.0012757982

0.0054267435

0.022996299

0.0110744992

0.0009707207

0.0058452799

0.019697275

0.0094596958

0.0008337325

0.0057302471

0.0165592562

0.0078128669

0.000458579

0.0057044639

0.0136848865

0.0065068179

0.0002559841

0.0055692522

0.0118703237

0.0056337584

0.0003236175

0.0058448599

0.0099907005

0.0048333546

-0.0000972686

0.005913212

0.008843259

0.0040129991

-0.0002274781

0.0055891275

0.0076358183

0.0031817518

-0.0003583728

0.0056866291

0.0064300178

0.0028176184

-0.0005092924

0.0058201198

0.0056945495

0.0026926273

-0.0004601122

0.0059322631

0.0053633689

0.0025610058

-0.0005242885

0.0057926523

0.0048543741

0.001971907

-0.0008198626

0.005901299

0.0042982087

0.0019222167

-0.0005559335

0.0059026107

0.0042067948

0.0016774102

-0.0002918439

0.0058515244

0.0041881329

0.0018605212

-0.000361373

0.0056446414

0.0036588192

0.001627

-0.0008881707

0.005611626

0.0031419352

0.0010933229

-0.0008768039

0.0058149775

0.0031494705

0.0013304779

-0.0006545251

0.0059100632

0.0036630481

0.0013135488

-0.0003961852

0.00583085

0.0030450292

0.0011256334

-0.00095404

0.0059312135

0.0029621928

0.0012714625

-0.0006043274

0.0059436001

0.0032079071

0.0010578749

-0.0005411454

0.0061487611

0.0032542502

0.0012540425

-0.000164293

0.0061611538

0.0028265032

0.0011245173

-0.0003450782

0.0059697386

0.0029498651

0.0008476084

-0.0005830264

0.0058907242

0.0029297192

0.0012810426

-0.0003815468

0.0057736333

0.0029189561

0.000905762

-0.0008405835

0.0059722317

0.0032688815

0.0009285379

-0.0005477639

0.006048216

0.0029759542

0.0010514915

-0.0005180833

0.0063055321

0.0030867662

0.0008608881

-0.0003480785

0.0055913562

0.002779268

0.0009108463

-0.000595228

0.0061876993

0.0029817405

0.0011457475

-0.0004707658

0.0063944729

0.0030362704

0.0014850515

-0.0002070384

0.0062225703

0.0026437659

0.0008384789

-0.0004199779

0.0061043138

0.0028867985

0.0007955569

-0.0002129375

0.0061894585

0.0034748702

0.0007809047

-0.0008616134

0.0063420413

0.002531192

0.0010081334

-0.0005438342

0.006474236

0.0030768856

0.0008773067

-0.0003816503

0.0060514705

0.0026810886

0.0007942083

-0.0003198328

0.0065214806

0.0030334292

0.0010330942

-0.0004354425

0.0065735658

0.002763585

0.001207056

0.0000541309

0.0065387976

0.0030639551

0.001069163

0.0000349734

0.0063608224

0.0028101669

0.0005377255

-0.0003723394

0.0066931597

0.0025632191

0.0011065581

0.0002176779

0.0066128573

0.0028570406

0.0012858976

0.000062079

0.0067295176

0.0027259425

0.0011874582

-0.0001089145

0.006789411

0.0029421765

0.0012391931

0.0000405911

0.0064367959

0.0029843468

0.0010828128

0.0002738045

0.005698143

0.0021576397

0.0003557671

-0.0002273228

0.007211132

0.0028728314

0.0011473048

0.0001389781

0.006281422

0.0024101096

0.0004331847

-0.0006103245

0.0064352457

0.0024468654

0.0004323814

-0.0002359124

0.0067762905

0.0029944079

0.0007776372

-0.0000936971

0.0072501651

0.0029630528

0.000952716

0.0000579884

0.0063014873

0.0019616617

0.0005521621

-0.0004978641

0.0070779235

0.0020814748

0.0005949822

-0.0007085419

0.0070485347

0.0025460075

0.0005887093

0.0000655223

0.006880397

0.0025289007

0.0007369253

-0.0000777545

0.0067300699

0.0023044925

0.0000847589

-0.0003339559

0.0065715946

0.0019305891

0.0001588395

-0.0011212259

0.0069476757

0.0024287996

0.0009035831

-0.0003761672

0.0076891966

0.0029523671

0.0008813272

-0.0001140386

0.0066509433

0.0019740653

0.000689165

-0.0003852711

0.0062568928

0.0018876371

0.000176138

-0.0004078231

0.0068921791

0.0020832175

0.0001509673

-0.0006529742

0.0065091564

0.002220236

0.0006102761

-0.0008149535

0.0071538645

0.0021520725

0.0004594082

-0.0005503492

0.006710405

0.0018008359

0.0006959061

-0.0003958748

0.0067334874

0.0018421937

0.0002357557

-0.0005179282

0.0071191029

0.0016704979

-0.0000676089

-0.0008623885

0.0066871662

0.0018098298

0.0001283875

-0.0009795584

0.0069605908

0.001956461

0.0006032512

-0.0007054923

0.0070273825

0.0015882349

0.0001328153

-0.0008145918

0.0074894694

0.001999602

0.0003402749

0.0000850437

0.0064644879

0.0015002729

0.0000864677

-0.0011115693

0.0069904211

0.0017833686

-0.0005354576

-0.0008849673

0.0070873168

0.001239297

-0.0004034266

-0.0003882712

0.0067464751

0.0016687829

-0.0001592213

-0.0006932421

0.0063916096

0.0009305095

-0.0004732482

-0.001465363

0.0072735134

0.0011379872

0.0004465294

0.0001687575

0.0074483869

0.0011821369

0.000260206

-0.0004542164

0.007149891

0.0008606028

-0.0007375376

-0.0007403286

0.0072166063

0.0010392958

-0.0004869009

-0.0003884781

0.0072568511

0.0009642607

-0.0010191763

-0.00122511

0.0064382409

0.0009065921

0.0001232865

-0.0007752137

0.0073088668

0.0018207735

0.0000598784

-0.0000570489

0.0072203963

0.0012238252

-0.0003937024

-0.0005653955

0.0071475487

0.0013026439

-0.0007416724

-0.0008927689

0.0073841107

0.0012023835

0.0000891863

-0.0004895383

0.0069052242

0.0010458088

-0.0009646815

-0.0011969212

0.0072223176

0.0010267891

-0.0005688597

-0.0004608363

0.0062258793

0.000951341

-0.0004638359

-0.0006722558

0.005964437

0.0002266649

-0.0011822582

-0.0007689088

0.0068462356

0.001437936

-0.0001560643

-0.0004506996

0.0070445095

0.0000436977

-0.0007037866

-0.0007150545

0.0072638262

0.0001245294

-0.0007171736

-0.0005460576

0.0064320932

0.0000430764

-0.0002446053

-0.001081824

0.0071122614

0.0013595861

-0.0001126929

-0.0006199405

0.0075494414

0.0009902827

0.0001169168

-0.0003490097

0.0069488068

0.001475415

0.0001344725

-0.0002615766

0.006690531

0.0006660646

-0.000643204

-0.0006915364

0.0063049281

0.0003770115

0.0004255667

0.0001857974

0.0074644508

0.0010660749

-0.0000281113

-0.0000814814

0.0083187725

0.0011221815

0.0011102434

0.0000405652

0.0067111147

0.0000743768

-0.000170762

-0.0005849394

0.0063071079

0.0001662974

-0.0003402154

-0.0008197076

0.0067987451

0.0012925442

-0.0002139725

-0.0005889722

0.0076478054

0.0012534194

0.0002136636

-0.0001791457

0.0071111824

0.0011909363

0.0004890282

0.0001527282

0.0066838013

0.000776574

-0.0000258852

-0.0006625893

0.0073030749

0.0009160344

-0.0004168746

-0.0006050511

0.0076204068

0.0012237732

0.0001810842

0.0002598952

0.0077395779

0.0009304576

-0.000197879

-0.000793973

0.0068267761

-0.0003762707

-0.0014471487

-0.0009316714

0.0073957224

0.0007789079

-0.0006009668

0.0006569389

0.0076335263

0.0014902204

-0.0002941203

0.0002394853

0.0081879413

0.0002233757

0.0007871286

0.00080606

0.0067921979

0.0009154637

-0.0002899813

0.0000355429

0.0073139477

0.0008153964

0.0012606366

0.0012093922

0.0073607308

0.0006295369

0.0000575483

-0.000262094

0.0070814495

0.0005516437

-0.0003098998

-0.0004577333

0.0075571854

0.0008047893

-0.0002264949

-0.0001808535

0.0066532562

0.0001499315

-0.0007645676

-0.0007855495

0.0063765836

0.0005169915

0.0007961015

-0.0009661795

0.0060943114

0.000118548

-0.0003110379

-0.0003266616

0.0060736258

0.0011510941

-0.0009035153

-0.0006177691

0.0061042611

0.0005787038

0.0003818305

-0.0004371492

0.0060338578

-0.0000986143

0.0003161309

-0.0007016157

0.0062491461

0.0013461353

-0.0007523191

0.000517147

0.0077264812

0.00085845

0.00102928

0.0010159173

0.0072704069

0.0001201534

-0.000319212

0.0006496281

0.0074872044

0.0005189612

-0.000513119

0.0005276175

0.0062703392

-0.0004442865

-0.0013486852

-0.0004051852

0.0061637005

-0.000003624

-0.0009105932

-0.0000403802

0.0072767246

0.0004457262

-0.0011562353

-0.0006166318

0.0062044254

-0.0012136488

-0.0006892104

-0.0010025961

0.0072027361

0.000302298

0.0005203607

0.0003830482

0.0068770042

0.0004183634

-0.0002064692

-0.0000124251

0.0075419345

0.0005322827

0.0001205418

0.0002772236

0.0068811332

-0.0003589418

-0.0003165735

-0.0008003293

0.0056334175

-0.0006136332

-0.0001835445

-0.0007229108

0.0060566156

0.000222262

-0.0015643674

-0.0006072225

0.0057069291

-0.0002771503

-0.0024162959

-0.0004844186

0.0058076582

-0.0014865175

-0.001075575

-0.0018737719

0.0072564823

-0.0005109989

-0.0021017136

0.0001288018

0.00757349

0.0012203986

-0.000130601

-0.0001775932

0.0060007866

-0.0008392917

-0.0021367471

-0.0017406003

0.007528475

-0.0001942049

-0.000140124

0.0003683582

0.007321029

0.0005172766

-0.0010343096

-0.0007836374

0.0055899927

-0.000546523

-0.0014338874

-0.0008609417

0.0050577382

-0.0018646992

-0.0023125419

-0.002317949

0.006661668

-0.0019281012

-0.0020813618

-0.0009399371

0.0057976102

-0.0013932236

-0.0023553336

-0.002474058

0.0063595879

-0.0012799851

-0.0016205407

-0.0009088883

0.005823845

-0.0007023393

-0.0020932639

-0.0012344025

0.0055073509

-0.0023986879

-0.0026361314

-0.0031502589

0.0067146374

-0.0017122386

-0.001889494

-0.0016873305

0.0072787781

0.0006048065

-0.0019951013

-0.0003897196

0.0055343024

-0.0013853795

-0.0039340351

-0.0021221158

0.0054130089

-0.0026353081

-0.0034461063

-0.0027088362

0.0068449993

-0.0000855706

-0.0006993931

0.000609991

0.00641943

-0.0003626145

-0.0017669491

-0.0016937252

0.0059720217

-0.0006301248

-0.0038509129

-0.0039248001

0.0058392193

-0.0013133816

-0.0007781595

-0.0018928961

0.0056040213

-0.0006867292

-0.0017749411

-0.0013269047

0.0058772112

-0.000007921

-0.0021475656

-0.0021986156

0.0043324842

-0.0011519496

-0.0036161351

-0.0021319042

0.0062730177

-0.0007499417

-0.0016013012

-0.0020398826

0.0070434571

-0.0016316817

-0.0012736875

-0.000633175

0.006261305

-0.000886104

-0.0026596987

-0.0011295913

0.0059828344

0.0010091712

-0.0011610372

-0.000310624

0.0049338024

-0.0007747486

-0.002482706

-0.0019179472



Sheet5

		350		350		350

		351		351		351

		352		352		352

		353		353		353

		354		354		354

		355		355		355

		356		356		356

		357		357		357

		358		358		358

		359		359		359

		360		360		360

		361		361		361

		362		362		362

		363		363		363

		364		364		364

		365		365		365

		366		366		366

		367		367		367

		368		368		368

		369		369		369

		370		370		370

		371		371		371

		372		372		372

		373		373		373

		374		374		374

		375		375		375

		376		376		376

		377		377		377

		378		378		378

		379		379		379

		380		380		380

		381		381		381

		382		382		382

		383		383		383

		384		384		384

		385		385		385

		386		386		386

		387		387		387

		388		388		388

		389		389		389

		390		390		390

		391		391		391

		392		392		392

		393		393		393

		394		394		394

		395		395		395

		396		396		396

		397		397		397

		398		398		398

		399		399		399

		400		400		400

		401		401		401

		402		402		402

		403		403		403

		404		404		404

		405		405		405

		406		406		406

		407		407		407

		408		408		408

		409		409		409

		410		410		410

		411		411		411

		412		412		412

		413		413		413

		414		414		414

		415		415		415

		416		416		416

		417		417		417

		418		418		418

		419		419		419

		420		420		420

		421		421		421

		422		422		422

		423		423		423

		424		424		424

		425		425		425

		426		426		426

		427		427		427

		428		428		428

		429		429		429

		430		430		430

		431		431		431

		432		432		432

		433		433		433

		434		434		434

		435		435		435

		436		436		436

		437		437		437

		438		438		438

		439		439		439

		440		440		440

		441		441		441

		442		442		442

		443		443		443

		444		444		444

		445		445		445

		446		446		446

		447		447		447

		448		448		448

		449		449		449

		450		450		450

		451		451		451

		452		452		452

		453		453		453

		454		454		454

		455		455		455

		456		456		456

		457		457		457

		458		458		458

		459		459		459

		460		460		460

		461		461		461

		462		462		462

		463		463		463

		464		464		464

		465		465		465

		466		466		466

		467		467		467

		468		468		468

		469		469		469

		470		470		470

		471		471		471

		472		472		472

		473		473		473

		474		474		474

		475		475		475

		476		476		476

		477		477		477

		478		478		478

		479		479		479

		480		480		480

		481		481		481

		482		482		482

		483		483		483

		484		484		484

		485		485		485

		486		486		486

		487		487		487

		488		488		488

		489		489		489

		490		490		490

		491		491		491

		492		492		492

		493		493		493

		494		494		494

		495		495		495

		496		496		496

		497		497		497

		498		498		498

		499		499		499

		500		500		500

		501		501		501

		502		502		502

		503		503		503

		504		504		504

		505		505		505

		506		506		506

		507		507		507

		508		508		508

		509		509		509

		510		510		510

		511		511		511

		512		512		512

		513		513		513

		514		514		514

		515		515		515

		516		516		516

		517		517		517

		518		518		518

		519		519		519

		520		520		520

		521		521		521

		522		522		522

		523		523		523

		524		524		524

		525		525		525

		526		526		526

		527		527		527

		528		528		528

		529		529		529

		530		530		530

		531		531		531

		532		532		532

		533		533		533

		534		534		534

		535		535		535

		536		536		536

		537		537		537

		538		538		538

		539		539		539

		540		540		540

		541		541		541

		542		542		542

		543		543		543

		544		544		544

		545		545		545

		546		546		546

		547		547		547

		548		548		548

		549		549		549

		550		550		550

		551		551		551

		552		552		552

		553		553		553

		554		554		554

		555		555		555

		556		556		556

		557		557		557

		558		558		558

		559		559		559

		560		560		560

		561		561		561

		562		562		562

		563		563		563

		564		564		564

		565		565		565

		566		566		566

		567		567		567

		568		568		568

		569		569		569

		570		570		570

		571		571		571

		572		572		572

		573		573		573

		574		574		574

		575		575		575

		576		576		576

		577		577		577

		578		578		578

		579		579		579

		580		580		580

		581		581		581

		582		582		582

		583		583		583

		584		584		584

		585		585		585

		586		586		586

		587		587		587

		588		588		588

		589		589		589

		590		590		590

		591		591		591

		592		592		592

		593		593		593

		594		594		594

		595		595		595

		596		596		596

		597		597		597

		598		598		598

		599		599		599

		600		600		600

		601		601		601

		602		602		602

		603		603		603

		604		604		604

		605		605		605

		606		606		606

		607		607		607

		608		608		608

		609		609		609

		610		610		610

		611		611		611

		612		612		612

		613		613		613

		614		614		614

		615		615		615

		616		616		616

		617		617		617

		618		618		618

		619		619		619

		620		620		620

		621		621		621

		622		622		622

		623		623		623

		624		624		624

		625		625		625

		626		626		626

		627		627		627

		628		628		628

		629		629		629

		630		630		630

		631		631		631

		632		632		632

		633		633		633

		634		634		634

		635		635		635

		636		636		636

		637		637		637

		638		638		638

		639		639		639

		640		640		640

		641		641		641

		642		642		642

		643		643		643

		644		644		644

		645		645		645

		646		646		646

		647		647		647

		648		648		648

		649		649		649

		650		650		650

		651		651		651

		652		652		652

		653		653		653

		654		654		654

		655		655		655

		656		656		656

		657		657		657

		658		658		658

		659		659		659

		660		660		660

		661		661		661

		662		662		662

		663		663		663

		664		664		664

		665		665		665

		666		666		666

		667		667		667

		668		668		668

		669		669		669

		670		670		670

		671		671		671

		672		672		672

		673		673		673

		674		674		674

		675		675		675

		676		676		676

		677		677		677

		678		678		678

		679		679		679

		680		680		680

		681		681		681

		682		682		682

		683		683		683

		684		684		684

		685		685		685

		686		686		686

		687		687		687

		688		688		688

		689		689		689

		690		690		690

		691		691		691

		692		692		692

		693		693		693

		694		694		694

		695		695		695

		696		696		696

		697		697		697

		698		698		698

		699		699		699

		700		700		700

		701		701		701

		702		702		702

		703		703		703

		704		704		704

		705		705		705

		706		706		706

		707		707		707

		708		708		708

		709		709		709

		710		710		710

		711		711		711

		712		712		712

		713		713		713

		714		714		714

		715		715		715

		716		716		716

		717		717		717

		718		718		718

		719		719		719

		720		720		720

		721		721		721

		722		722		722

		723		723		723

		724		724		724

		725		725		725

		726		726		726

		727		727		727

		728		728		728

		729		729		729

		730		730		730

		731		731		731

		732		732		732

		733		733		733

		734		734		734

		735		735		735

		736		736		736

		737		737		737

		738		738		738

		739		739		739

		740		740		740

		741		741		741

		742		742		742

		743		743		743

		744		744		744

		745		745		745

		746		746		746

		747		747		747

		748		748		748

		749		749		749

		750		750		750

		751		751		751

		752		752		752

		753		753		753

		754		754		754

		755		755		755

		756		756		756

		757		757		757

		758		758		758

		759		759		759

		760		760		760

		761		761		761

		762		762		762

		763		763		763

		764		764		764

		765		765		765

		766		766		766

		767		767		767

		768		768		768

		769		769		769

		770		770		770

		771		771		771

		772		772		772

		773		773		773

		774		774		774

		775		775		775

		776		776		776

		777		777		777

		778		778		778

		779		779		779

		780		780		780

		781		781		781

		782		782		782

		783		783		783

		784		784		784

		785		785		785

		786		786		786

		787		787		787

		788		788		788

		789		789		789

		790		790		790

		791		791		791

		792		792		792

		793		793		793

		794		794		794

		795		795		795

		796		796		796

		797		797		797

		798		798		798

		799		799		799

		800		800		800

		801		801		801

		802		802		802

		803		803		803

		804		804		804

		805		805		805

		806		806		806

		807		807		807

		808		808		808

		809		809		809

		810		810		810

		811		811		811

		812		812		812

		813		813		813

		814		814		814

		815		815		815

		816		816		816

		817		817		817

		818		818		818

		819		819		819

		820		820		820

		821		821		821

		822		822		822

		823		823		823

		824		824		824

		825		825		825

		826		826		826

		827		827		827

		828		828		828

		829		829		829

		830		830		830

		831		831		831

		832		832		832

		833		833		833

		834		834		834

		835		835		835

		836		836		836

		837		837		837

		838		838		838

		839		839		839

		840		840		840

		841		841		841

		842		842		842

		843		843		843

		844		844		844

		845		845		845

		846		846		846

		847		847		847

		848		848		848

		849		849		849

		850		850		850



Dil1

Dil_2

Dil_3

0.2928262279

0.2589933041

-0.3440597129

0.2984823032

0.2721746112

-0.2809552886

0.3032027547

0.267654755

-0.2986825864

0.3153510185

0.2855516599

-0.3029736958

0.3174678333

0.2839864601

-0.2908726895

0.3250410437

0.293339126

-0.2572436086

0.3367073574

0.3045122298

-0.2146206551

0.3468921377

0.3152565505

-0.2233474913

0.3544052731

0.3206304093

-0.2341757043

0.3660176345

0.3329389639

-0.1851564718

0.3734801093

0.3337979953

-0.1820280635

0.3807249642

0.3474821193

-0.2049287631

0.3895840443

0.3525945122

-0.1938351548

0.4031407968

0.3629870476

-0.1277933787

0.4090125827

0.3701831428

-0.1529198679

0.4204922522

0.3854606412

-0.1351786263

0.4370696116

0.4035427637

-0.0671079401

0.4483138199

0.4102432357

-0.0915720656

0.4551356147

0.4135565715

-0.1003280235

0.4747897391

0.4534500466

0.0102651075

0.4857585895

0.4656310488

0.013707156

0.4987156447

0.4755777164

0.0815198094

0.5126743569

0.4966134632

0.1152067445

0.5235104826

0.50464228

0.1167867281

0.5313675765

0.5071981934

0.117808459

0.5426371314

0.5201571994

0.1345687436

0.5483222073

0.528137593

0.1515983137

0.5478319619

0.5260695431

0.1158381933

0.5547076433

0.5319363276

0.1344521721

0.5585769381

0.5375709198

0.1405717501

0.5594522235

0.5399571184

0.1561415188

0.5630760171

0.5452167129

0.1756566937

0.560343902

0.5378219287

0.1488807447

0.5580586796

0.5331631634

0.1433023876

0.5602084122

0.5424968795

0.1705661527

0.56005304

0.5427991424

0.1767014617

0.5588439019

0.5443652746

0.2164858496

0.5559186249

0.5376825201

0.185597934

0.5506312294

0.5271289953

0.0893602503

0.5485546731

0.5240233404

0.128368651

0.5495032008

0.5310580801

0.154033954

0.5476940032

0.532868194

0.146703011

0.5465349002

0.5324681852

0.1616780564

0.548935763

0.53400148

0.1478420429

0.5542908375

0.5419628573

0.1556397101

0.5584107689

0.5486601988

0.1769077141

0.5663898563

0.5521777849

0.1726093592

0.5794410565

0.5680793583

0.213555257

0.593368398

0.5838449201

0.2548399163

0.6127169131

0.6050573421

0.3009620604

0.6355685449

0.6290431494

0.2892870286

0.653940651

0.642022737

0.3020314947

0.6810550117

0.6723238942

0.3490161643

0.7093239987

0.6998414752

0.3912150883

0.7317731112

0.7212685051

0.4104470826

0.7567692747

0.7477500131

0.4259269759

0.7782989259

0.7699802578

0.463205141

0.7967304158

0.7900215941

0.5034549643

0.8113869533

0.8038230464

0.5027579416

0.8244204437

0.8164248386

0.5075551632

0.8333827213

0.824222196

0.5287536667

0.8381272357

0.8318985948

0.5391145721

0.8445578678

0.8366153748

0.5497952016

0.8431383188

0.8353679574

0.546087862

0.8424371029

0.8341388571

0.5648601588

0.8292506762

0.820302969

0.5491385395

0.8245043842

0.8265647668

0.5537025935

0.8267250219

0.8172264548

0.5760861911

0.816796474

0.8095193501

0.5124813442

0.8176791988

0.8116431832

0.5302578527

0.8228844972

0.8222656797

0.5516065657

0.8348534905

0.8312283277

0.5632398511

0.8517126267

0.8502950027

0.593183815

0.8743165188

0.8750236122

0.6321456976

0.8999075075

0.8961168075

0.6439998479

0.9344449739

0.9360643344

0.7087450519

0.966487662

0.9705517464

0.749326124

1.0019690793

1.0058231237

0.7762005916

1.0389074409

1.0428406719

0.840402435

1.0664917375

1.0684646803

0.8952541385

1.0924342227

1.099481131

0.9006968864

1.1074728708

1.1092304751

0.9329864332

1.1075256705

1.1089603169

0.9014496359

1.0972329202

1.1009748612

0.9280391967

1.0716001257

1.0726912988

0.8861208759

1.0299196855

1.031996374

0.8307378714

0.9721896907

0.9714166394

0.7753728589

0.9037880604

0.9032407007

0.6900283847

0.8218846507

0.8218458554

0.6072358185

0.7372488626

0.7356345855

0.5340596263

0.6466793042

0.6477649996

0.408845934

0.5619239585

0.559812621

0.3186939535

0.4794355

0.4855430343

0.2420978086

0.4040525223

0.4021505093

0.1522980716

0.3348858815

0.3395275361

0.0824432586

0.2739228279

0.274464296

0.0069044284

0.2242297406

0.2298104462

-0.0387415345

0.1797295747

0.1830061871

-0.1162764583

0.1439705607

0.1502890248

-0.1464967464

0.1167444768

0.1233321422

-0.1643412966

0.0909937624

0.0948322817

-0.1992773314

0.0718778612

0.0771704165

-0.2144000335

0.0581230298

0.0724603555

-0.1949387131

0.0483158376

0.056375929

-0.2073451957

0.0391815479

0.0487735015

-0.2276618931

0.0302151124

0.0387286591

-0.2494427077

0.0251299025

0.0324303253

-0.2358430103

0.0210115183

0.0322025915

-0.2251604601

0.0181326329

0.0277004028

-0.2352729048

0.0159042575

0.0253564894

-0.2479640997

0.0130936941

0.022907443

-0.2554803368

0.0097902806

0.0206981633

-0.242565408

0.0104480295

0.0219716718

-0.2376716128

0.0076008913

0.0183034724

-0.2600013071

0.0075243424

0.0193988778

-0.2356601476

0.0074198354

0.0181522348

-0.2573941341

0.006159098

0.018270071

-0.2426729471

0.0044583776

0.0163686805

-0.2525288345

0.0062654036

0.0187112742

-0.2435709704

0.007223892

0.0173376396

-0.2310453268

0.00549819

0.0169670181

-0.2402852774

0.0057098512

0.0202953612

-0.2380426871

0.0049846614

0.0208967856

-0.2436032345

0.0044034177

0.0170319569

-0.2379243771

0.0051390332

0.0197951379

-0.2228794542

0.0071173307

0.0200392133

-0.2220240528

0.0046715849

0.0178033782

-0.221986391

0.0038584381

0.0142859886

-0.2584746554

0.0063180994

0.0192105005

-0.2347942427

0.0066553694

0.0213678433

-0.2162886661

0.0067085287

0.0203162425

-0.2368003344

0.0064988744

0.019832723

-0.2390752619

0.0089505999

0.0231474387

-0.1990781786

0.0088384879

0.0223225725

-0.2357247003

0.0097373086

0.0254372859

-0.2242997096

0.008657374

0.0215219284

-0.2334388561

0.0087442616

0.0210736024

-0.2283880678

0.0094146716

0.0185892563

-0.2694610771

0.0105824372

0.0170212883

-0.2596680161

0.0116415759

0.0205408457

-0.2442646621

0.0134019491

0.0222650157

-0.2610387446

0.0139620595

0.0298824069

-0.2174616278

0.0122530598

0.0233414838

-0.2228633101

0.0149621982

0.0280264798

-0.2136090386

0.0150471623

0.0289155968

-0.219985016

0.0131285808

0.0281305305

-0.2212977566

0.0152003321

0.0268885208

-0.2081847204

0.014832001

0.0322560374

-0.1872186186

0.0163060002

0.0381444568

-0.164077348

0.0150611709

0.0359568723

-0.1776716478

0.0181130982

0.0397162297

-0.1576450239

0.0168433305

0.0358834088

-0.1569554169

0.0173186981

0.039128697

-0.1315723949

0.0182902929

0.0371133709

-0.1392965863

0.0177660886

0.0345522914

-0.1442710663

0.0172582681

0.0321784235

-0.1548595592

0.0167335155

0.0299167886

-0.1879939067

0.0169584339

0.0325358288

-0.1840634412

0.0171787695

0.0298968106

-0.1811449689

0.0187058477

0.0343877199

-0.1817103698

0.0158530373

0.0312536155

-0.1776824187

0.0187104439

0.0331465699

-0.1715428669

0.0190653558

0.0352273541

-0.1538412322

0.0198483931

0.03555514

-0.1616693765

0.0182673101

0.0349177103

-0.1655802454

0.021229097

0.0333892049

-0.1823242276

0.0212700386

0.0396706413

-0.1656879789

0.0234902565

0.0375500636

-0.1793894896

0.0251727265

0.040531313

-0.1351516746

0.0253903114

0.0443864473

-0.148167274

0.0270945771

0.0410821951

-0.1540190377

0.028306217

0.0411808402

-0.1679610829

0.0304523513

0.0453528576

-0.150107172

0.0319508147

0.0480641986

-0.1317502975

0.0330426316

0.0474909323

-0.1318904534

0.0346766713

0.0518643654

-0.1348498212

0.034666977

0.0482528131

-0.1368765187

0.0355388513

0.047202686

-0.1398247873

0.0356037282

0.0502230505

-0.1221485118

0.0371979722

0.0520358328

-0.1098319531

0.0381773891

0.0498559789

-0.1382886898

0.0381457386

0.0487832844

-0.1205848547

0.0390565342

0.0494036877

-0.1336585514

0.038573613

0.0503106263

-0.1280090161

0.0374656725

0.0442933514

-0.1355505516

0.0381626049

0.0532483147

-0.1096431943

0.037310687

0.0506781849

-0.1473050768

0.0392929336

0.0496603414

-0.1250977286

0.0408235187

0.0498340859

-0.1457746201

0.0407909218

0.0494907889

-0.1758244819

0.0410149441

0.0500744501

-0.1492180596

0.0385622933

0.049583481

-0.137959909

0.0401180152

0.0520185899

-0.1504843612

0.041785592

0.0540522371

-0.1385042908

0.0408921812

0.0561321055

-0.1301923664

0.040506812

0.055072058

-0.1327691108

0.0428818302

0.0598755896

-0.112194142

0.040927321

0.0592748687

-0.1158396521

0.0406952173

0.0557750125

-0.1238954205

0.0420802067

0.0626311993

-0.1151332323

0.0411118172

0.0619406879

-0.0978087685

0.0422816351

0.0609642823

-0.1045786081

0.0441926392

0.0637179437

-0.0931151351

0.0452597081

0.0616131851

-0.0963144114

0.0449104663

0.060119996

-0.0942912879

0.0453446526

0.0618641772

-0.1323217039

0.0460462553

0.0618856371

-0.100435909

0.0480615103

0.0598965792

-0.1116278781

0.0507325463

0.0611714012

-0.1186221262

0.0540665353

0.0654503313

-0.1154567856

0.0555496528

0.0674776966

-0.1146640647

0.0576409562

0.0704209551

-0.1173926841

0.0618490661

0.0719446169

-0.1132403689

0.0634054217

0.079473914

-0.0919875123

0.0683176361

0.0811172188

-0.0905091535

0.0708359868

0.085095035

-0.1018276384

0.0748936951

0.0860805699

-0.0826633211

0.0789092275

0.0928080677

-0.0656341705

0.0800296685

0.0961203543

-0.0789449796

0.0886510572

0.094805106

-0.0803157871

0.0898312441

0.1018191891

-0.0891170777

0.091555242

0.100887456

-0.07924721

0.0960819226

0.1041886442

-0.070222684

0.1002300562

0.1117112557

-0.0634693108

0.0999182666

0.1084415976

-0.0601975252

0.1032492313

0.110944067

-0.0640847673

0.1058190614

0.1139912075

-0.0672968872

0.107042881

0.1123888328

-0.0670377649

0.1086762725

0.1161097455

-0.0777144616

0.1106630042

0.1193878044

-0.0663143803

0.1133556097

0.1184060244

-0.0589557014

0.1142929122

0.1211796679

-0.0449206931

0.1145144329

0.1212276167

-0.0652292759

0.1160307686

0.1183510486

-0.0561587668

0.1147001389

0.1159457838

-0.0575572554

0.1138452049

0.1190235636

-0.056477342

0.1134305307

0.1163488815

-0.0636258722

0.111279646

0.111213093

-0.0703576306

0.1110764441

0.1133984961

-0.0529188564

0.1070146539

0.1067948542

-0.0804614986

0.1055639481

0.1045606448

-0.0764569251

0.1056831601

0.1108097955

-0.0544416423

0.104017435

0.1043260879

-0.0627134735

0.1043833827

0.1082059917

-0.0545010404

0.1058310518

0.1028021513

-0.0445264175

0.1061047554

0.105505949

-0.0364133298

0.1077297446

0.1087283744

-0.0606942463

0.1114451815

0.112717547

-0.0325885494

0.1143822266

0.1098276908

-0.0502457571

0.1209327126

0.1179713383

-0.0418096145

0.1255847279

0.1229315089

-0.0507425876

0.1295630713

0.1283225608

-0.028255569

0.1363637494

0.1291578218

-0.0395139749

0.1425084324

0.1386546087

-0.0198531584

0.1503265352

0.1485739259

0.0135773638

0.1563282807

0.1536920136

0.0056040228

0.1608361652

0.1585229415

0.0095053504

0.1679934094

0.1665532903

0.0140965324

0.174983974

0.1736363034

0.0257360983

0.1790409664

0.1738208882

0.0121199413

0.1864317955

0.1850957511

0.0391534283

0.1883794454

0.1829493214

0.0427113091

0.1922139955

0.1862225428

0.006944982

0.1965449302

0.1913977629

0.0312134597

0.1979003881

0.1983876092

0.0406225419

0.2003898165

0.1938872356

0.0449084056

0.2024991259

0.1941496675

0.0458608762

0.2026992664

0.1961658302

0.0500200478

0.2027233787

0.1962987281

0.042362272

0.2007358982

0.1983994819

0.0459366404

0.2007380707

0.1934848457

0.0547777172

0.2008286936

0.1964985826

0.0470298443

0.1965615372

0.1921254951

0.0482042525

0.1914089239

0.1836160748

0.0219226206

0.1883799228

0.1816775592

0.0379657266

0.1853336233

0.177191123

0.0324523811

0.1791912364

0.1728303078

0.0174820828

0.1747972311

0.1671533868

0.0226068589

0.1694112466

0.1636872878

0.0208814084

0.1623625656

0.1575158925

-0.0003892552

0.1592281286

0.1500932

-0.004497856

0.1554844976

0.1540437156

0.0094620904

0.1482875381

0.1430834713

0.0132691113

0.1448624038

0.1407049493

-0.0343605225

0.1427420339

0.1341585728

-0.020226032

0.141065594

0.1349991707

-0.0331233881

0.1407231991

0.1328509927

-0.0058817219

0.1403770186

0.1346497334

-0.0147731346

0.141886993

0.1373561704

-0.0066385058

0.144410162

0.1404466804

0.0060176612

0.1497081077

0.1424197152

0.0068206171

0.1539729762

0.1497661925

-0.0094979538

0.1659884659

0.1578035518

0.0303992515

0.1764719427

0.1711878273

0.0584396754

0.1858480963

0.1796878893

0.0472273802

0.2022102554

0.1989944497

0.0514651577

0.2223498521

0.2158284555

0.0724108908

0.2455197994

0.2416434597

0.1365479689

0.2707045492

0.2649039659

0.1416753693

0.301976631

0.3010680112

0.1858313568

0.3409148172

0.3363287313

0.2316143625

0.3839418274

0.3816147487

0.2778869594

0.4303202361

0.4313994747

0.3171531999

0.4856212892

0.4858228838

0.405681955

0.5433398944

0.5433014597

0.4287048547

0.6077776567

0.6091822517

0.560672674

0.6742670894

0.6749633328

0.6268063312

0.7382618104

0.7428306809

0.6931573521

0.8042067819

0.8055762112

0.7645537392

0.8623330417

0.868318944

0.8493253362

0.9162971037

0.9124218199

0.8979285386

0.955500015

0.9519888874

0.9479721766

0.985071364

0.9827769422

0.9605304491

1

1

1

0.9984227088

0.9939251017

0.9864253213

0.9859392758

0.9805432739

0.9804142184

0.9548674664

0.9506208446

0.9498243876

0.914462658

0.9081638961

0.8819146211

0.8639602365

0.8572978709

0.8477252087

0.8018303112

0.7971200434

0.7571262225

0.731645785

0.7225972497

0.6736898804

0.6575277741

0.6492442761

0.6329187247

0.5881326835

0.5787979687

0.5084982396

0.5176099264

0.5144298567

0.4795690285

0.451038298

0.4467715218

0.4068575181

0.3904411262

0.3833261841

0.3437127576

0.3327684084

0.326454965

0.260932251

0.2837689824

0.2739100251

0.1817329936

0.239999609

0.2359327189

0.1332267173

0.2032025638

0.1980087316

0.10136864

0.1740513457

0.1691365292

0.0870634953

0.1463228203

0.1396917225

0.0478876468

0.1209239812

0.1163399562

0.0267314398

0.1048899307

0.1007299136

0.0337941338

0.0882809865

0.0864189339

-0.0101573845

0.0781418303

0.0717512232

-0.0237546635

0.0674725034

0.0568887702

-0.0374234936

0.0568176686

0.0503781742

-0.0531834554

0.0503188389

0.0481433694

-0.048047758

0.0473924222

0.0457900165

-0.0547494344

0.0428947834

0.0352571066

-0.0856151091

0.0379803306

0.0343686598

-0.0580540068

0.0371725679

0.0299915925

-0.0304761372

0.0370076659

0.0332655613

-0.0377368018

0.0323304829

0.0290902739

-0.0927482609

0.0277631323

0.0195482858

-0.0915612765

0.0278297169

0.0237885463

-0.0683495451

0.0323678508

0.0234858588

-0.0413721002

0.0269068405

0.0201259877

-0.0996267459

0.0261748711

0.0227333696

-0.0631075911

0.0283460806

0.0189144862

-0.0565097399

0.0287555826

0.0224219045

-0.0171564932

0.0249758758

0.0201060326

-0.0360351894

0.0260659398

0.0151549841

-0.0608832177

0.0258879242

0.022904658

-0.039843473

0.0257928184

0.0161947538

-0.0877789172

0.0288848695

0.0166019801

-0.0572008847

0.0262964717

0.0188003529

-0.0541014523

0.0272756417

0.0153924216

-0.0363485067

0.0245584899

0.0162856589

-0.0621573847

0.0263476013

0.0204856222

-0.0491602729

0.0268294445

0.0265522758

-0.0216202262

0.0233611507

0.0149917509

-0.0438566839

0.0255086634

0.0142243188

-0.022236251

0.0307050513

0.0139623427

-0.0899749896

0.0223664126

0.0180251235

-0.0567905196

0.0271883333

0.0156859819

-0.0398542773

0.0236909463

0.014200207

-0.03339891

0.0268043384

0.0184714159

-0.045471591

0.0244199099

0.0215817998

0.0056526861

0.0270740755

0.0191163139

0.0036521332

0.0248315228

0.0096143702

-0.038881979

0.0226494134

0.0197849283

0.0227312671

0.0252457136

0.0229914636

0.0064826708

0.0240872895

0.0212313966

-0.0113735233

0.0259980015

0.0221564024

0.0042387725

0.0263706315

0.0193603685

0.0285923581

0.0190655862

0.0063610095

-0.0237384539

0.0253852462

0.020513466

0.0145129516

0.0212964901

0.0077452123

-0.0637338489

0.0216212761

0.0077308501

-0.0246354286

0.0264595343

0.013903921

-0.0097844289

0.026182471

0.0170342767

0.0060555106

0.0173338629

0.0098724924

-0.05199004

0.0183925686

0.0106381033

-0.0739903153

0.0224973268

0.0105259453

0.0068422456

0.022346166

0.0131760046

-0.008119603

0.020363224

0.0015154643

-0.0348737311

0.0170592952

0.0028400021

-0.1170853235

0.021461641

0.0161557949

-0.0392817039

0.0260880483

0.0157578665

-0.0119086142

0.0174434649

0.0123220644

-0.0402323874

0.0166797576

0.0031492939

-0.0425874039

0.018407968

0.0026992504

-0.0681875982

0.0196187068

0.0109115545

-0.0851024748

0.0190163926

0.0082140813

-0.0574708631

0.0159127554

0.012442593

-0.0413396932

0.0162782052

0.0042152418

-0.0540852533

0.0147610476

-0.0012088267

-0.0900559235

0.0159922284

0.0022955299

-0.1022915361

0.0172879075

0.0107859512

-0.0736718616

0.0140341455

0.0023746984

-0.0850647036

0.0176691147

0.0060840131

0.0088807915

0.0132568853

0.0015460174

-0.1160769287

0.0157584088

-0.0095738211

-0.0924137482

0.0109508197

-0.0072131466

-0.0405456774

0.014745893

-0.0028468291

-0.0723926192

0.0082222757

-0.008461536

-0.1530222403

0.0100556149

0.0079838129

0.0176227021

0.0104457355

0.0046524059

-0.0474320826

0.007604558

-0.0131869521

-0.0773096764

0.0091835461

-0.0087056421

-0.0405672831

0.0085205119

-0.0182225677

-0.1279335467

0.0080109339

0.0022043253

-0.0809525971

0.0160889297

0.0010706074

-0.005957401

0.0108141057

-0.0070392808

-0.0590420877

0.0115105732

-0.0132608797

-0.0932284365

0.0106246409

0.0015946246

-0.0511206095

0.0092410968

-0.0172482165

-0.1249898977

0.0090730327

-0.010171042

-0.0481233672

0.0084063497

-0.0082932465

-0.0702010935

0.0020028828

-0.0211384236

-0.0802941967

0.0127060566

-0.0027903837

-0.0470648346

0.0003861268

-0.0125834941

-0.0746704035

0.0011003807

-0.0128228497

-0.0570227024

0.0003806366

-0.004373469

-0.1129707309

0.0120137327

-0.002014915

-0.0647380072

0.0087504512

0.0020904368

-0.0364457429

0.0130372328

0.0024043286

-0.0273154406

0.0058855574

-0.0115002944

-0.0722144976

0.0033313929

0.0076090045

0.0194021151

0.009420175

-0.0005026208

-0.0085087887

0.0099159506

0.0198508193

0.0042360691

0.0006572165

-0.0030531742

-0.0610829844

0.0014694565

-0.00608295

-0.0855989223

0.0114213288

-0.0038257637

-0.0615041084

0.0110756102

0.0038202405

-0.018707501

0.0105234895

0.0087436773

0.0159488193

0.0068620534

-0.0004628196

-0.0691916592

0.0080943693

-0.007453593

-0.06318317

0.0108136469

0.0032377314

0.0271398629

0.0082218174

-0.0035380177

-0.0829115631

-0.0033248468

-0.0258745863

-0.0972908764

0.0068826764

-0.0107451065

0.0686016104

0.0131680586

-0.0052587832

0.0250085362

0.0019738181

0.0140736234

0.0841737617

0.0080893262

-0.0051847796

0.0037116078

0.0072051005

0.0225398052

0.1262921955

0.0055627867

0.0010289468

-0.0273694715

0.0048744977

-0.0055409151

-0.0477993306

0.0071113726

-0.0040496615

-0.0188858368

0.0013248423

-0.0136702395

-0.082031927

0.0045683006

0.0142340559

-0.1008944273

0.0010475278

-0.0055612651

-0.0341120161

0.0101714311

-0.0161545825

-0.0645112585

0.0051136098

0.0068270155

-0.0456498188

-0.0008713875

0.0056523264

-0.07326704

0.0118948767

-0.0134512396

0.0540036812

0.0075855356

0.0184032186

0.1060883498

0.0010617134

-0.0057074146

0.067838169

0.0045857054

-0.0091744149

0.0550970765

-0.0039258559

-0.0241140884

-0.0423119368

-0.000032023

-0.0162811338

-0.0042167531

0.0039385771

-0.0206731419

-0.0643924926

-0.0107241845

-0.0123228751

-0.1046972707

0.0026712003

0.0093038938

0.0400002593

0.0036967909

-0.0036916069

-0.0012975043

0.004703418

0.0021552512

0.0289494035

-0.0031717231

-0.0056602401

-0.0835753186

-0.0054222568

-0.0032817207

-0.0754908055

0.0019639773

-0.0279704216

-0.0634099187

-0.0024489874

-0.043202648

-0.0505859805

-0.0131353382

-0.0192309593

-0.1956708204

-0.0045153476

-0.0375780094

0.0134502781

0.0107838275

-0.0023351059

-0.0185453779

-0.0074162468

-0.038204399

-0.1817642244

-0.0017160557

-0.0025053757

0.0384662345

0.0045708193

-0.0184931465

-0.0818322514

-0.0048292496

-0.0256374776

-0.0899048448

-0.0164770711

-0.0413475577

-0.2420545303

-0.0170373112

-0.0372141266

-0.0981540279

-0.0123109644

-0.0421126607

-0.258356394

-0.0113103525

-0.0289747835

-0.0949117226

-0.0062060923

-0.037426932

-0.1289039273

-0.021195564

-0.0471332398

-0.3289694695

-0.0151298817

-0.0337835875

-0.1762014644

0.0053442616

-0.0356718148

-0.0406969203

-0.0122416507

-0.0703393697

-0.2216044241

-0.0232864148

-0.0616153498

-0.2828733895

-0.0007561287

-0.0125049393

0.0636990197

-0.0032041767

-0.0315925206

-0.1768692378

-0.005567982

-0.0688531702

-0.4098518367

-0.0116054548

-0.0139132585

-0.1976678918

-0.0060681563

-0.0317354157

-0.1385635796

-0.0000699928

-0.0383978301

-0.2295930018

-0.0101789909

-0.0646554133

-0.2226265928

-0.0066267217

-0.0286307871

-0.2130171238

-0.014418055

-0.0227731506

-0.0661200355

-0.007829895

-0.047554617

-0.1179588873

0.008917356

-0.0207589985

-0.0324372793

-0.0068459236

-0.0443900399

-0.2002838784



								In vitro		In vivo				118.2321279285		0.054543378

		Blank		Scale		Fo		Chl		Chl				chl		Fo

		0		1		5		591.22		10304.43				127.59		1.0759493671

		0		3.16		5		187.13		3260.43				63.79		0.545

		0		10		5		59.17		1029.83				12.76		0.1060126582

		0.05		31.6		5		18.58		322.17				6.38		0.05

		0.1		100		5		5.85		100.31				1.28		0.0107594937

		0.4		316		5		1.78		29.32				0.64		0.0054

		1.3		1000		5		0.49		6.95				0.13		0.0010126582

		4.2		3160		5		0.08		-0.16				0.06		0.0005063291





		



Chl (ug/L)

Fo



		





		






_1211720987.xls
Chart2

		3.16

		10

		31.6

		100

		316

		1000

		3160

		3160



scale

Fo/Fa

7.1578947368

7.6542222222

7.9761904762

7.3488372093

7.1729957806

7.4191304348

10.6666666667

16



cal

		March 30 2003										10 cm cells

		serial dilution of Sigma Chl (stored in CCMP freezer)																																										137.3658635236

		ID		Vol sln		Vol Acoh		dilution		OD663		OD725		chl ug/l				ODest chl		Fest chl		Fest pheo		Fo/Fa		Fo-Fa		Scale		Ro		Blank		Scale		Ra		Blank		Fo		Fa		F/chl		F/pheo

										10 cm						127.59						0.31		7.45		137.260499495																		0.0083711833		0.0010913543		7.6704545455

		1		3		57		1		0.1131693348		0.0012391931		127.67		127.67		127.59		12.71		0.61		7.16		0.93		3.16		3.4		0		31.6		4.8		0.05		1.0759		0.1503		0.0084331921		0.0011781665		7.1578947368

		2		25		25		0.5		0.0559293479		4.06E-05		63.75		127.50		63.79		6.50		-0.19		7.65		0.47		10		5.45		0		31.6		2.3		0.05		0.5450		0.0712		0.0085433196		0.0011161578		7.6542222222

		3		10		40		0.1										12.76		1.27		-0.10		7.98		0.09		31.6		3.4		0.05		316		4.6		0.4		0.1060		0.0133		0.0083091745		0.0010417473		7.9761904762

		4		5		5		0.05										6.38		0.59		0.01		7.35		0.04		100		5.1		0.1		316		2.55		0.4		0.0500		0.0068		0.0078379079		0.0010665508		7.3488372093

		5		4		16		0.01										1.28		0.13		0.01		7.17		0.01		316		3.8		0.4		1000		2.8		1.3		0.0108		0.0015		0.0084331921		0.0011756862		7.1729957806

		6		10		10		0.005										0.64		0.06		0.00		7.42		0.00		1000		6.7		1.3		3160		6.5		4.2		0.0054		0.0007		0.0084649405		0.0011409613		7.4191304348

		7		4		16		0.001										0.13		0.01		-0.00		10.67		0.00		3160		7.4		4.2		3160		4.5		4.2		0.0010		0.0001		0.0079371219		0.0007441052		10.6666666667

		8		10		10		0.0005										0.06		0.01		-0.00		16.00		0.00		3160		5.8		4.2		3160		4.3		4.2		0.0005		0.0000		0.0079371219		0.0004960701		16

																						7.20		0.0063042766





cal

		



&A

Page &P

chl

Fo-Fa



OD

		



&A

Page &P

scale

Fo/Fa



Finvivo

		538.2000684385

		269.1000342192

		134.5500171096

		67.2750085548

		33.6375042774

		20.6997544089

		16.0210427959

		8.010521398

		4.005260699

		2.0026303495

		1.0013151747

		0.5006575874



&A

Page &P

meas Chl

est Chl

544.5089383462

294.6692737414

153.3827995341

77.4836662936

38.8052144689

19.8713121115

15.4967332587

7.7800572904

3.6051296023

1.7777827042

0.9022331204

0.5002053942



Sheet4

		zero_acoh				Chl1		Chl2		Chl3

		Wavelength (nm)		Abs		Abs		Abs		Abs

		350		0.0063626086		0.0331389494		0.0144853266		-0.0032947653		0.2928262279		0.2589933041		-0.3440597129

		351		0.0066823554		0.0337790437		0.0152225485		-0.0026904682		0.2984823032		0.2721746112		-0.2809552886

		352		0.0063110213		0.0343132541		0.0149697559		-0.0028602274		0.3032027547		0.267654755		-0.2986825864

		353		0.0065864171		0.035688065		0.0159707181		-0.0029013196		0.3153510185		0.2855516599		-0.3029736958

		354		0.0064050336		0.0359276235		0.0158831775		-0.0027854387		0.3174678333		0.2839864601		-0.2908726895

		355		0.0063150665		0.0367846787		0.016406266		-0.0024634018		0.3250410437		0.293339126		-0.2572436086

		356		0.0064474363		0.0381049477		0.0170311704		-0.0020552382		0.3367073574		0.3045122298		-0.2146206551

		357		0.0065289699		0.0392575525		0.0176320933		-0.0021388077		0.3468921377		0.3152565505		-0.2233474913

		358		0.0064026695		0.040107809		0.0179326497		-0.0022425002		0.3544052731		0.3206304093		-0.2341757043

		359		0.0067532319		0.0414219722		0.0186210591		-0.001773085		0.3660176345		0.3329389639		-0.1851564718

		360		0.0069384696		0.0422664955		0.0186691042		-0.0017431269		0.3734801093		0.3337979953		-0.1820280635

		361		0.0067374045		0.0430863909		0.0194344483		-0.0019624273		0.3807249642		0.3474821193		-0.2049287631

		362		0.0063128602		0.0440889671		0.0197203811		-0.0018561933		0.3895840443		0.3525945122		-0.1938351548

		363		0.0066746534		0.0456231758		0.0203016289		-0.0012237678		0.4031407968		0.3629870476		-0.1277933787

		364		0.0066737072		0.0462876819		0.0207041018		-0.0014643826		0.4090125827		0.3701831428		-0.1529198679

		365		0.0066943429		0.0475868285		0.0215585623		-0.0012944899		0.4204922522		0.3854606412		-0.1351786263

		366		0.0072340039		0.0494628772		0.0225698836		-0.0006426353		0.4370696116		0.4035427637		-0.0671079401

		367		0.0068470505		0.0507353768		0.0229446366		-0.0008769073		0.4483138199		0.4102432357		-0.0915720656

		368		0.006724759		0.0515073948		0.0231299494		-0.0009607556		0.4551356147		0.4135565715		-0.1003280235

		369		0.0069450191		0.0537316389		0.0253611654		9.83E-05		0.4747897391		0.4534500466		0.0102651075

		370		0.0067010201		0.0549729765		0.0260424409		0.0001312617		0.4857585895		0.4656310488		0.013707156

		371		0.007078818		0.0564393178		0.0265987515		0.0007806454		0.4987156447		0.4755777164		0.0815198094

		372		0.0072587198		0.0580190159		0.0277752671		0.0011032363		0.5126743569		0.4966134632		0.1152067445

		373		0.0072689592		0.0592453331		0.0282243136		0.0011183665		0.5235104826		0.50464228		0.1167867281

		374		0.0069234511		0.0601345152		0.0283672642		0.0011281507		0.5313675765		0.5071981934		0.117808459

		375		0.0071638124		0.0614098832		0.0290920529		0.0012886496		0.5426371314		0.5201571994		0.1345687436

		376		0.0073622051		0.0620532595		0.0295383912		0.0014517272		0.5483222073		0.528137593		0.1515983137

		377		0.0067888061		0.0619977787		0.0294227265		0.0011092832		0.5478319619		0.5260695431		0.1158381933

		378		0.0070389584		0.062775895		0.0297508519		0.0012875333		0.5547076433		0.5319363276		0.1344521721

		379		0.0070695304		0.0632137805		0.030065991		0.0013461353		0.5585769381		0.5375709198		0.1405717501

		380		0.0071684439		0.063312836		0.0301994495		0.0014952336		0.5594522235		0.5399571184		0.1561415188

		381		0.0071280757		0.0637229383		0.0304936152		0.0016821137		0.5630760171		0.5452167129		0.1756566937

		382		0.0069729015		0.0634137467		0.0300800297		0.0014257034		0.560343902		0.5378219287		0.1488807447

		383		0.0068736644		0.0631551295		0.029819468		0.0013722843		0.5580586796		0.5331631634		0.1433023876

		384		0.0072780144		0.0633984134		0.0303414967		0.001633366		0.5602084122		0.5424968795		0.1705661527

		385		0.0071341284		0.06338083		0.0303584021		0.0016921186		0.56005304		0.5427991424		0.1767014617

		386		0.007377265		0.0632439926		0.0304459948		0.0020730996		0.5588439019		0.5443652746		0.2164858496

		387		0.007266643		0.062912941		0.0300722327		0.0017773125		0.5559186249		0.5376825201		0.185597934

		388		0.0069699553		0.06231457		0.029481981		0.0008557266		0.5506312294		0.5271289953		0.0893602503

		389		0.0071860766		0.0620795675		0.0293082837		0.0012292766		0.5485546731		0.5240233404		0.128368651

		390		0.0070838174		0.0621869117		0.0297017321		0.0014750513		0.5495032008		0.5310580801		0.154033954

		391		0.0069564614		0.061982166		0.0298029706		0.0014048491		0.5476940032		0.532868194		0.146703011

		392		0.0071095773		0.0618509911		0.0297805984		0.0015482523		0.5465349002		0.5324681852		0.1616780564

		393		0.0070247515		0.0621226951		0.0298663545		0.0014157566		0.548935763		0.53400148		0.1478420429

		394		0.0071298652		0.0627287254		0.0303116292		0.0014904282		0.5542908375		0.5419628573		0.1556397101

		395		0.0068829479		0.0631949753		0.0306862071		0.0016940937		0.5584107689		0.5486601988		0.1769077141

		396		0.0069075911		0.0640979633		0.0308829434		0.001652932		0.5663898563		0.5521777849		0.1726093592

		397		0.0070196479		0.0655749589		0.0317723081		0.0020450358		0.5794410565		0.5680793583		0.213555257

		398		0.0070487452		0.0671511069		0.0326540656		0.0024403837		0.593368398		0.5838449201		0.2548399163

		399		0.0071524438		0.0693407655		0.0338404626		0.0028820559		0.6127169131		0.6050573421		0.3009620604

		400		0.0071725757		0.0719268694		0.0351819731		0.0027702542		0.6355685449		0.6290431494		0.2892870286

		401		0.0067546782		0.0740060285		0.035907913		0.002892297		0.653940651		0.642022737		0.3020314947

		402		0.0070010489		0.0770745426		0.037602637		0.0033422289		0.6810550117		0.6723238942		0.3490161643

		403		0.007079476		0.0802737251		0.0391416773		0.0037463319		0.7093239987		0.6998414752		0.3912150883

		404		0.0070278826		0.0828142762		0.0403400771		0.0039305002		0.7317731112		0.7212685051		0.4104470826

		405		0.0069124834		0.0856430754		0.0418211706		0.0040787379		0.7567692747		0.7477500131		0.4259269759

		406		0.0070482455		0.0880795717		0.0430644937		0.004435719		0.7782989259		0.7699802578		0.463205141

		407		0.007000654		0.0901654512		0.0441853926		0.0048211571		0.7967304158		0.7900215941		0.5034549643

		408		0.0067726881		0.0918241218		0.0449572988		0.0048144823		0.8113869533		0.8038230464		0.5027579416

		409		0.0069482545		0.0932991132		0.0456621088		0.0048604212		0.8244204437		0.8164248386		0.5075551632

		410		0.0069893426		0.0943133682		0.0460982099		0.0050634211		0.8333827213		0.824222196		0.5287536667

		411		0.0069949455		0.0948503017		0.0465275459		0.0051626386		0.8381272357		0.8318985948		0.5391145721

		412		0.0071093403		0.0955780521		0.0467913523		0.0052649179		0.8445578678		0.8366153748		0.5497952016

		413		0.0069454927		0.0954174027		0.0467215851		0.0052294158		0.8431383188		0.8353679574		0.546087862

		414		0.0072105266		0.0953380465		0.0466528423		0.005409182		0.8424371029		0.8341388571		0.5648601588

		415		0.0067665619		0.0938457474		0.0458790101		0.0052586296		0.8292506762		0.820302969		0.5491385395

		416		0.0071643386		0.0933086127		0.0462292284		0.0053023356		0.8245043842		0.8265647668		0.5537025935

		417		0.0073758699		0.0935599208		0.0457069427		0.0055166841		0.8267250219		0.8172264548		0.5760861911

		418		0.0068793711		0.0924363136		0.0452758893		0.004907595		0.816796474		0.8095193501		0.5124813442

		419		0.0067896475		0.092536211		0.0453946739		0.0050778254		0.8176791988		0.8116431832		0.5302578527

		420		0.0068943095		0.0931252912		0.0459887832		0.0052822637		0.8228844972		0.8222656797		0.5516065657

		421		0.0070326966		0.0944798142		0.0464900583		0.0053936658		0.8348534905		0.8312283277		0.5632398511

		422		0.00701094		0.0963877514		0.047556445		0.0056804135		0.8517126267		0.8502950027		0.593183815

		423		0.0070712143		0.0989458188		0.0489395		0.006053518		0.8743165188		0.8750236122		0.6321456976

		424		0.0070973942		0.101841934		0.0501192287		0.0061670351		0.8999075075		0.8961168075		0.6439998479

		425		0.00717626		0.1057505161		0.0523534678		0.0067870445		0.9344449739		0.9360643344		0.7087450519

		426		0.0070877904		0.1093767658		0.0542823262		0.0071756546		0.966487662		0.9705517464		0.749326124

		427		0.0070246202		0.1133921742		0.0562550314		0.0074330084		1.0019690793		1.0058231237		0.7762005916

		428		0.00723074		0.117572464		0.0583253987		0.0080478145		1.0389074409		1.0428406719		0.840402435

		429		0.0071469699		0.1206941605		0.0597585328		0.0085730823		1.0664917375		1.0684646803		0.8952541385

		430		0.0072196596		0.1236300543		0.0614932626		0.0086252028		1.0924342227		1.099481131		0.9006968864

		431		0.0071124453		0.1253319681		0.0620385371		0.0089344122		1.1074728708		1.1092304751		0.9329864332

		432		0.0069485963		0.1253379434		0.0620234273		0.0086324112		1.1075256705		1.1089603169		0.9014496359

		433		0.0071322336		0.1241731197		0.061576806		0.0088870367		1.0972329202		1.1009748612		0.9280391967

		434		0.0071783918		0.1212722734		0.0599949248		0.0084856208		1.0716001257		1.0726912988		0.8861208759

		435		0.0070760818		0.1165553257		0.0577188842		0.0079552652		1.0299196855		1.031996374		0.8307378714

		436		0.0073061027		0.1100220606		0.0543306992		0.0074250819		0.9721896907		0.9714166394		0.7753728589

		437		0.0071101035		0.1022810936		0.0505176634		0.006607811		0.9037880604		0.9032407007		0.6900283847

		438		0.0069205314		0.0930121392		0.0459653027		0.0058149775		0.8218846507		0.8218458554		0.6072358185

		439		0.0072960462		0.0834339634		0.0411435626		0.0051142317		0.7372488626		0.7356345855		0.5340596263

		440		0.0071857609		0.0731842667		0.036229074		0.0039151674		0.6466793042		0.6477649996		0.408845934

		441		0.006989737		0.0635925606		0.0313099548		0.0030518591		0.5619239585		0.559812621		0.3186939535

		442		0.0072748293		0.0542573966		0.0271561053		0.0023183634		0.4794355		0.4855430343		0.2420978086

		443		0.0070522181		0.0457263552		0.0224920157		0.0014584282		0.4040525223		0.4021505093		0.1522980716

		444		0.0073972233		0.0378988124		0.0189895537		0.0007894885		0.3348858815		0.3395275361		0.0824432586

		445		0.007051534		0.0309996642		0.0153506091		6.61E-05		0.2739228279		0.274464296		0.0069044284

		446		0.0070892903		0.0253759306		0.0128531484		-0.0003709945		0.2242297406		0.2298104462		-0.0387415345

		447		0.0071311281		0.0203398764		0.0102354167		-0.0011134801		0.1797295747		0.1830061871		-0.1162764583

		448		0.0071483646		0.0162930526		0.0084055671		-0.0014028739		0.1439705607		0.1502890248		-0.1464967464

		449		0.0072281607		0.0132118948		0.0068978863		-0.0015737559		0.1167444768		0.1233321422		-0.1643412966

		450		0.0068754		0.0102977036		0.0053039077		-0.0019083084		0.0909937624		0.0948322817		-0.1992773314

		451		0.0069497013		0.0081343697		0.0043160911		-0.0020531255		0.0718778612		0.0771704165		-0.2144000335

		452		0.0070714774		0.0065777446		0.0040526604		-0.0018667611		0.0581230298		0.0724603555		-0.1949387131

		453		0.0072846483		0.0054678712		0.0031530689		-0.0019855674		0.0483158376		0.056375929		-0.2073451957

		454		0.0071790498		0.0044341497		0.0027278701		-0.002180123		0.0391815479		0.0487735015		-0.2276618931

		455		0.0070928163		0.0034194242		0.0021660686		-0.0023886992		0.0302151124		0.0387286591		-0.2494427077

		456		0.0070518763		0.0028439343		0.0018138069		-0.0022584666		0.0251299025		0.0324303253		-0.2358430103

		457		0.0073525426		0.0023778595		0.0018010699		-0.0021561689		0.0210115183		0.0322025915		-0.2251604601

		458		0.0072014462		0.002052058		0.0015492655		-0.0022530071		0.0181326329		0.0277004028		-0.2352729048

		459		0.0070701879		0.0017998742		0.0014181719		-0.0023745398		0.0159042575		0.0253564894		-0.2479640997

		460		0.007039879		0.0014818046		0.0012811983		-0.0024465164		0.0130936941		0.022907443		-0.2554803368

		461		0.0072447956		0.0011079595		0.0011576348		-0.0023228412		0.0097902806		0.0206981633		-0.242565408

		462		0.0071706017		0.0011823965		0.0012288613		-0.0022759775		0.0104480295		0.0219716718		-0.2376716128

		463		0.0070822388		0.0008601878		0.0010237013		-0.0024898099		0.0076008913		0.0183034724		-0.2600013071

		464		0.0073059183		0.0008515248		0.0010849666		-0.0022567154		0.0075243424		0.0193988778		-0.2356601476

		465		0.0072318721		0.0008396978		0.0010152427		-0.0024648432		0.0074198354		0.0181522348		-0.2573941341

		466		0.0071305232		0.000697021		0.0010218332		-0.002323871		0.006159098		0.018270071		-0.2426729471

		467		0.0071157082		0.0005045516		0.0009154896		-0.0024182524		0.0044583776		0.0163686805		-0.2525288345

		468		0.0070989733		0.0007090516		0.0010465094		-0.0023324706		0.0062654036		0.0187112742		-0.2435709704

		469		0.0071602594		0.000817523		0.0009696829		-0.0022125232		0.007223892		0.0173376396		-0.2310453268

		470		0.0071256021		0.0006222265		0.0009489543		-0.0023010063		0.00549819		0.0169670181		-0.2402852774

		471		0.0072764088		0.0006461801		0.0011351063		-0.002279531		0.0057098512		0.0202953612		-0.2380426871

		472		0.0071547329		0.0005641108		0.0011687436		-0.0023327796		0.0049846614		0.0208967856		-0.2436032345

		473		0.0070751607		0.0004983319		0.0009525862		-0.002278398		0.0044034177		0.0170319569		-0.2379243771

		474		0.0072319247		0.000581581		0.0011071292		-0.0021343257		0.0051390332		0.0197951379		-0.2228794542

		475		0.0071856817		0.0008054636		0.0011207801		-0.0021261342		0.0071173307		0.0200392133		-0.2220240528

		476		0.0071145506		0.0005286802		0.0009957313		-0.0021257736		0.0046715849		0.0178033782		-0.221986391

		477		0.0070083621		0.0004366569		0.000799006		-0.0024751904		0.0038584381		0.0142859886		-0.2584746554

		478		0.0072153956		0.0007150151		0.0010744308		-0.0022484234		0.0063180994		0.0192105005		-0.2347942427

		479		0.0070241992		0.0007531837		0.0011950895		-0.0020712113		0.0066553694		0.0213678433		-0.2162886661

		480		0.0068443683		0.0007591997		0.0011362742		-0.002267634		0.0067085287		0.0203162425		-0.2368003344

		481		0.007088948		0.0007354733		0.0011092313		-0.0022894191		0.0064988744		0.019832723		-0.2390752619

		482		0.007093579		0.0010129334		0.0012946211		-0.0019064012		0.0089505999		0.0231474387		-0.1990781786

		483		0.0070432206		0.0010002458		0.0012484869		-0.0022573336		0.0088384879		0.0223225725		-0.2357247003

		484		0.0070732664		0.0011019647		0.0014226908		-0.0021479262		0.0097373086		0.0254372859		-0.2242997096

		485		0.006913404		0.0009797493		0.0012037074		-0.002235444		0.008657374		0.0215219284		-0.2334388561

		486		0.0069320258		0.0009895823		0.0011786328		-0.002187077		0.0087442616		0.0210736024		-0.2283880678

		487		0.0070231208		0.0010654521		0.001039685		-0.002580398		0.0094146716		0.0185892563		-0.2694610771

		488		0.0066966298		0.0011976074		0.0009519896		-0.0024866182		0.0105824372		0.0170212883		-0.2596680161

		489		0.0068685361		0.0013174694		0.0011488361		-0.0023391135		0.0116415759		0.0205408457		-0.2442646621

		490		0.0068562021		0.0015166897		0.0012452678		-0.0024997445		0.0134019491		0.0222650157		-0.2610387446

		491		0.0070241732		0.001580077		0.0016713035		-0.0020824438		0.0139620595		0.0298824069		-0.2174616278

		492		0.0065723308		0.0013866706		0.001305474		-0.0021341711		0.0122530598		0.0233414838		-0.2228633101

		493		0.0065321755		0.001693262		0.0015675027		-0.0020455508		0.0149621982		0.0280264798		-0.2136090386

		494		0.006686877		0.0017028773		0.0016172305		-0.0021066081		0.0150471623		0.0289155968		-0.219985016

		495		0.006327989		0.0014857528		0.0015733222		-0.0021191791		0.0131285808		0.0281305305		-0.2212977566

		496		0.0063599558		0.0017202115		0.0015038574		-0.0019936068		0.0152003321		0.0268885208		-0.2081847204

		497		0.0066238176		0.0016785277		0.0018040591		-0.0017928324		0.014832001		0.0322560374		-0.1872186186

		498		0.0068143904		0.0018453392		0.0021333946		-0.0015712283		0.0163060002		0.0381444568		-0.164077348

		499		0.0067692962		0.0017044627		0.0020110444		-0.0017014092		0.0150611709		0.0359568723		-0.1776716478

		500		0.0071256286		0.0020498473		0.0022213028		-0.0015096314		0.0181130982		0.0397162297		-0.1576450239

		501		0.0069961557		0.0019061485		0.0020069357		-0.0015030276		0.0168433305		0.0358834088		-0.1569554169

		502		0.0071245497		0.0019599455		0.0021884425		-0.0012599562		0.0173186981		0.039128697		-0.1315723949

		503		0.0068884711		0.0020699003		0.0020757266		-0.0013339241		0.0182902929		0.0371133709		-0.1392965863

		504		0.0069667725		0.0020105764		0.0019324871		-0.0013815605		0.0177660886		0.0345522914		-0.1442710663

		505		0.0068237521		0.0019531067		0.0017997182		-0.0014829574		0.0172582681		0.0321784235		-0.1548595592

		506		0.0068230154		0.0018937208		0.0016732265		-0.0018002567		0.0167335155		0.0299167886		-0.1879939067

		507		0.0067385351		0.0019191747		0.0018197077		-0.001762618		0.0169584339		0.0325358288		-0.1840634412

		508		0.0066319918		0.0019441099		0.0016721091		-0.0017346703		0.0171787695		0.0298968106		-0.1811449689

		509		0.0069110105		0.0021169283		0.0019232827		-0.0017400846		0.0187058477		0.0343877199		-0.1817103698

		510		0.0067201057		0.0017940777		0.0017479943		-0.0017015124		0.0158530373		0.0312536155		-0.1776824187

		511		0.0066745221		0.0021174485		0.001853866		-0.0016427191		0.0187104439		0.0331465699		-0.1715428669

		512		0.0067631439		0.0021576136		0.0019702429		-0.0014732057		0.0190653558		0.0352273541		-0.1538412322

		513		0.0069570662		0.0022462294		0.0019885758		-0.0015481692		0.0198483931		0.03555514		-0.1616693765

		514		0.0065745381		0.0020672993		0.0019529248		-0.0015856202		0.0182673101		0.0349177103		-0.1655802454

		515		0.0066250791		0.0024024828		0.0018674365		-0.001745963		0.021229097		0.0333892049		-0.1823242276

		516		0.0067924345		0.0024071161		0.0022187531		-0.0015866519		0.0212700386		0.0396706413		-0.1656879789

		517		0.0065992167		0.0026583767		0.0021001506		-0.0017178595		0.0234902565		0.0375500636		-0.1793894896

		518		0.0068949931		0.0028487807		0.0022668899		-0.0012942319		0.0251727265		0.040531313		-0.1351516746

		519		0.0068016378		0.0028734046		0.0024825051		-0.0014188711		0.0253903114		0.0443864473		-0.148167274

		520		0.0066106231		0.0030662753		0.0022977004		-0.0014749084		0.0270945771		0.0410821951		-0.1540190377

		521		0.0066502336		0.0032033958		0.0023032175		-0.0016084195		0.028306217		0.0411808402		-0.1679610829

		522		0.0066421377		0.0034462723		0.0025365558		-0.0014374478		0.0304523513		0.0453528576		-0.150107172

		523		0.0065820022		0.0036158524		0.0026881993		-0.0012616598		0.0319508147		0.0480641986		-0.1317502975

		524		0.0067594368		0.0037394126		0.0026561369		-0.0012630019		0.0330426316		0.0474909323		-0.1318904534

		525		0.0066507855		0.0039243358		0.0029007401		-0.0012913413		0.0346766713		0.0518643654		-0.1348498212

		526		0.0066321497		0.0039232387		0.0026987484		-0.0013107492		0.034666977		0.0482528131		-0.1368765187

		527		0.006399333		0.0040219082		0.0026400154		-0.0013389823		0.0355388513		0.047202686		-0.1398247873

		528		0.0065650772		0.0040292502		0.0028089425		-0.0011697117		0.0356037282		0.0502230505		-0.1221485118

		529		0.0064888713		0.0042096698		0.0029103302		-0.0010517666		0.0371979722		0.0520358328		-0.1098319531

		530		0.0063934224		0.0043205097		0.0027884124		-0.0013242724		0.0381773891		0.0498559789		-0.1382886898

		531		0.0063913208		0.0043169279		0.0027284173		-0.0011547379		0.0381457386		0.0487832844		-0.1205848547

		532		0.0063524432		0.004420002		0.002763116		-0.0012799335		0.0390565342		0.0494036877		-0.1336585514

		533		0.0061668251		0.0043653501		0.0028138405		-0.0012258327		0.038573613		0.0503106263		-0.1280090161

		534		0.0060198153		0.0042399652		0.0024772983		-0.0012980516		0.0374656725		0.0442933514		-0.1355505516

		535		0.0062376438		0.0043188366		0.0029781435		-0.001049959		0.0381626049		0.0532483147		-0.1096431943

		536		0.0060296319		0.0042224256		0.0028343978		-0.0014106146		0.037310687		0.0506781849		-0.1473050768

		537		0.0060608154		0.0044467552		0.0027774705		-0.0011979538		0.0392929336		0.0496603414		-0.1250977286

		538		0.0064643831		0.0046199705		0.0027871879		-0.0013959587		0.0408235187		0.0498340859		-0.1457746201

		539		0.0062465197		0.0046162815		0.0027679875		-0.0016837205		0.0407909218		0.0494907889		-0.1758244819

		540		0.0062750401		0.0046416339		0.0028006313		-0.0014289336		0.0410149441		0.0500744501		-0.1492180596

		541		0.0063723279		0.0043640691		0.0027731718		-0.0013211239		0.0385622933		0.049583481		-0.137959909

		542		0.0063972315		0.0045401291		0.0029093658		-0.0014410599		0.0401180152		0.0520185899		-0.1504843612

		543		0.0065267622		0.0047288476		0.0030231064		-0.001326337		0.041785592		0.0540522371		-0.1385042908

		544		0.006621636		0.0046277409		0.0031394321		-0.0012467408		0.0408921812		0.0561321055		-0.1301923664

		545		0.0065295738		0.004584129		0.0030801443		-0.0012714161		0.040506812		0.055072058		-0.1327691108

		546		0.0067503401		0.0048529082		0.0033488027		-0.0010743872		0.0428818302		0.0598755896		-0.112194142

		547		0.0066690543		0.0046317177		0.0033152048		-0.0011092972		0.040927321		0.0592748687		-0.1158396521

		548		0.0064558438		0.0046054507		0.0031194601		-0.0011864404		0.0406952173		0.0557750125		-0.1238954205

		549		0.0066660312		0.004762189		0.0035029221		-0.0011025324		0.0420802067		0.0626311993		-0.1151332323

		550		0.0066539138		0.004652597		0.0034643023		-0.0009366308		0.0411118172		0.0619406879		-0.0978087685

		551		0.0065748538		0.0047849845		0.0034096926		-0.0010014598		0.0422816351		0.0609642823		-0.1045786081

		552		0.0065837102		0.0050012516		0.003563703		-0.0008916839		0.0441926392		0.0637179437		-0.0931151351

		553		0.006489607		0.0051220111		0.0034459853		-0.0009223206		0.0452597081		0.0616131851		-0.0963144114

		554		0.0064041666		0.0050824876		0.0033624722		-0.0009029469		0.0449104663		0.060119996		-0.0942912879

		555		0.0061520431		0.0051316242		0.0034600231		-0.0012671317		0.0453446526		0.0618641772		-0.1323217039

		556		0.0064222151		0.0052110241		0.0034612233		-0.0009617887		0.0460462553		0.0618856371		-0.100435909

		557		0.0062488043		0.0054390891		0.0033499766		-0.0010689646		0.0480615103		0.0598965792		-0.1116278781

		558		0.0062399283		0.0057413685		0.0034212766		-0.0011359425		0.0507325463		0.0611714012		-0.1186221262

		559		0.0061940798		0.0061186738		0.0036605943		-0.0011056308		0.0540665353		0.0654503313		-0.1154567856

		560		0.0061704484		0.0062865173		0.0037739836		-0.0010980396		0.0555496528		0.0674776966		-0.1146640647

		561		0.0060493182		0.0065231887		0.0039385981		-0.0011241692		0.0576409562		0.0704209551		-0.1173926841

		562		0.0061451904		0.0069994177		0.0040238155		-0.001084406		0.0618490661		0.0719446169		-0.1132403689

		563		0.0057864608		0.0071755494		0.0044449242		-0.0008808856		0.0634054217		0.079473914		-0.0919875123

		564		0.0061539859		0.0077314614		0.0045368331		-0.0008667287		0.0683176361		0.0811172188		-0.0905091535

		565		0.005997532		0.0080164615		0.0047593098		-0.0009751161		0.0708359868		0.085095035		-0.1018276384

		566		0.0060569569		0.0084756697		0.0048144301		-0.0007915958		0.0748936951		0.0860805699		-0.0826633211

		567		0.006177275		0.0089301048		0.0051906947		-0.0006285222		0.0789092275		0.0928080677		-0.0656341705

		568		0.0058336048		0.0090569044		0.0053759487		-0.0007559885		0.0800296685		0.0961203543		-0.0789449796

		569		0.0058061364		0.0100325812		0.0053023878		-0.0007691155		0.0886510572		0.094805106		-0.0803157871

		570		0.0056898552		0.0101661421		0.0056946808		-0.0008533979		0.0898312441		0.1018191891		-0.0891170777

		571		0.0057484247		0.0103612458		0.0056425696		-0.0007588827		0.091555242		0.100887456		-0.07924721

		572		0.0059117163		0.0108735273		0.0058272029		-0.0006724625		0.0960819226		0.1041886442		-0.070222684

		573		0.0057182079		0.0113429688		0.0062479377		-0.0006077912		0.1002300562		0.1117112557		-0.0634693108

		574		0.0057008183		0.0113076838		0.0060650678		-0.0005764602		0.0999182666		0.1084415976		-0.0601975252

		575		0.0058653252		0.0116846468		0.0062050293		-0.0006136849		0.1032492313		0.110944067		-0.0640847673

		576		0.0058822227		0.0119754728		0.0063754539		-0.0006444447		0.1058190614		0.1139912075		-0.0672968872

		577		0.0057519921		0.0121139716		0.0062858341		-0.0006419633		0.107042881		0.1123888328		-0.0670377649

		578		0.0059043425		0.0122988215		0.0064939423		-0.0007442049		0.1086762725		0.1161097455		-0.0777144616

		579		0.0058376449		0.0125236586		0.006677282		-0.000635036		0.1106630042		0.1193878044		-0.0663143803

		580		0.0060197893		0.0128283789		0.0066223717		-0.0005645683		0.1133556097		0.1184060244		-0.0589557014

		581		0.0060512871		0.0129344529		0.0067774998		-0.000430167		0.1142929122		0.1211796679		-0.0449206931

		582		0.0059544384		0.0129595222		0.0067801815		-0.0006246449		0.1145144329		0.1212276167		-0.0652292759

		583		0.0060747806		0.0131311249		0.006619297		-0.0005377844		0.1160307686		0.1183510486		-0.0561587668

		584		0.0058824853		0.0129805384		0.0064847721		-0.0005511765		0.1147001389		0.1159457838		-0.0575572554

		585		0.0059282482		0.0128837861		0.0066569103		-0.0005408351		0.1138452049		0.1190235636		-0.056477342

		586		0.005922344		0.0128368577		0.0065073171		-0.0006092905		0.1134305307		0.1163488815		-0.0636258722

		587		0.0059157573		0.0125934435		0.0062200758		-0.0006737548		0.111279646		0.111213093		-0.0703576306

		588		0.0060184244		0.0125704473		0.0063423039		-0.0005067586		0.1110764441		0.1133984961		-0.0529188564

		589		0.0056754826		0.0121107772		0.0059729666		-0.0007705109		0.1070146539		0.1067948542		-0.0804614986

		590		0.0055845389		0.0119466018		0.0058480087		-0.0007321625		0.1055639481		0.1045606448		-0.0764569251

		591		0.0058773947		0.0119600929		0.0061975196		-0.000521341		0.1056831601		0.1108097955		-0.0544416423

		592		0.0056746174		0.0117715839		0.0058348901		-0.0006005532		0.104017435		0.1043260879		-0.0627134735

		593		0.0056793904		0.011812998		0.0060518906		-0.0005219098		0.1043833827		0.1082059917		-0.0545010404

		594		0.0058547254		0.0119768297		0.0057496573		-0.0004263914		0.1058310518		0.1028021513		-0.0445264175

		595		0.00567029		0.0120078046		0.0059008789		-0.0003486993		0.1061047554		0.105505949		-0.0364133298

		596		0.0056547644		0.0121917035		0.0060811071		-0.0005812168		0.1077297446		0.1087283744		-0.0606942463

		597		0.0056610848		0.0126121771		0.0063042189		-0.0003120726		0.1114451815		0.112717547		-0.0325885494

		598		0.0054750275		0.0129445605		0.0061425911		-0.0004811606		0.1143822266		0.1098276908		-0.0502457571

		599		0.0054219468		0.0136858746		0.00659806		-0.0004003749		0.1209327126		0.1179713383		-0.0418096145

		600		0.0053899176		0.0142123401		0.0068754791		-0.0004859183		0.1255847279		0.1229315089		-0.0507425876

		601		0.0054444103		0.0146625666		0.0071769971		-0.0002705794		0.1295630713		0.1283225608		-0.028255569

		602		0.0054828133		0.0154321948		0.0072237127		-0.0003783915		0.1363637494		0.1291578218		-0.0395139749

		603		0.0055087404		0.0161275845		0.0077548618		-0.0001901167		0.1425084324		0.1386546087		-0.0198531584

		604		0.0054706237		0.017012354		0.0083096428		0.0001300188		0.1503265352		0.1485739259		0.0135773638

		605		0.0055588693		0.0176915675		0.0085958941		5.37E-05		0.1563282807		0.1536920136		0.0056040228

		606		0.0055543333		0.0182017218		0.0088660847		9.10E-05		0.1608361652		0.1585229415		0.0095053504

		607		0.0055473591		0.0190117024		0.0093152169		0.0001349904		0.1679934094		0.1665532903		0.0140965324

		608		0.0055509508		0.0198028199		0.0097113652		0.0002464526		0.174983974		0.1736363034		0.0257360983

		609		0.0054692342		0.0202619471		0.0097216889		0.0001160623		0.1790409664		0.1738208882		0.0121199413

		610		0.005637587		0.0210983623		0.0103522846		0.0003749389		0.1864317955		0.1850957511		0.0391534283

		611		0.0053571574		0.0213187765		0.0102322362		0.0004090096		0.1883794454		0.1829493214		0.0427113091

		612		0.0052323765		0.02175273		0.0104153054		6.65E-05		0.1922139955		0.1862225428		0.006944982

		613		0.0054254327		0.022242859		0.0107047521		0.0002989046		0.1965449302		0.1913977629		0.0312134597

		614		0.0054379883		0.0223962553		0.0110956896		0.0003890073		0.1979003881		0.1983876092		0.0406225419

		615		0.0053262864		0.0226779822		0.0108439866		0.0004300493		0.2003898165		0.1938872356		0.0449084056

		616		0.0055094482		0.0229166914		0.0108586643		0.0004391703		0.2024991259		0.1941496675		0.0458608762

		617		0.0055114147		0.0229393411		0.010971427		0.0004789992		0.2026992664		0.1961658302		0.0500200478

		618		0.0052438783		0.0229420699		0.0109788598		0.0004056672		0.2027233787		0.1962987281		0.042362272

		619		0.0051972703		0.0227171481		0.0110963536		0.0004398959		0.2007358982		0.1983994819		0.0459366404

		620		0.0054638078		0.0227173939		0.0108214812		0.0005245593		0.2007380707		0.1934848457		0.0547777172

		621		0.0054097064		0.0227276497		0.0109900376		0.0004503645		0.2008286936		0.1964985826		0.0470298443

		622		0.0054363371		0.0222447384		0.0107454537		0.0004616108		0.1965615372		0.1921254951		0.0482042525

		623		0.0054697846		0.0216616206		0.0102695273		0.0002099342		0.1914089239		0.1836160748		0.0219226206

		624		0.0053207045		0.0213188306		0.0101611074		0.0003635653		0.1883799228		0.1816775592		0.0379657266

		625		0.0052827089		0.0209740829		0.009910184		0.0003107687		0.1853336233		0.177191123		0.0324523811

		626		0.0052697915		0.020278953		0.0096662864		0.0001674109		0.1791912364		0.1728303078		0.0174820828

		627		0.0053277276		0.0197816864		0.0093487799		0.0002164865		0.1747972311		0.1671533868		0.0226068589

		628		0.0052110241		0.0191721581		0.0091549233		0.0001999634		0.1694112466		0.1636872878		0.0208814084

		629		0.00525842		0.0183744635		0.0088097611		-3.73E-06		0.1623625656		0.1575158925		-0.0003892552

		630		0.0052906484		0.0180197414		0.0083946148		-4.31E-05		0.1592281286		0.1500932		-0.004497856

		631		0.0055256765		0.0175960772		0.0086155646		9.06E-05		0.1554844976		0.1540437156		0.0094620904

		632		0.0053404374		0.0167816021		0.0080025652		0.0001270669		0.1482875381		0.1430834713		0.0132691113

		633		0.005270394		0.0163939819		0.0078695361		-0.0003290413		0.1448624038		0.1407049493		-0.0343605225

		634		0.0051209894		0.016154021		0.0075034015		-0.0001936874		0.1427420339		0.1341585728		-0.020226032

		635		0.0053580222		0.0159642994		0.0075504156		-0.0003171943		0.141065594		0.1349991707		-0.0331233881

		636		0.0053817406		0.0159255508		0.0074302694		-5.63E-05		0.1407231991		0.1328509927		-0.0058817219

		637		0.0053559258		0.0158863738		0.0075308718		-0.0001414697		0.1403770186		0.1346497334		-0.0147731346

		638		0.005306161		0.0160572566		0.007682241		-6.36E-05		0.141886993		0.1373561704		-0.0066385058

		639		0.005404071		0.016342802		0.0078550912		5.76E-05		0.144410162		0.1404466804		0.0060176612

		640		0.0053970469		0.016942367		0.0079654418		6.53E-05		0.1497081077		0.1424197152		0.0068206171

		641		0.0052469959		0.0174250193		0.0083763255		-9.10E-05		0.1539729762		0.1497661925		-0.0094979538

		642		0.0054860641		0.0187848043		0.0088258497		0.0002911076		0.1659884659		0.1578035518		0.0303992515

		643		0.0056288545		0.0199712124		0.0095744235		0.0005596267		0.1764719427		0.1711878273		0.0584396754

		644		0.0051608048		0.0210323054		0.0100498265		0.0004522562		0.1858480963		0.1796878893		0.0472273802

		645		0.005377023		0.0228840001		0.0111296298		0.0004928377		0.2022102554		0.1989944497		0.0514651577

		646		0.0052400269		0.0251631848		0.0120711448		0.0006934171		0.2223498521		0.2158284555		0.0724108908

		647		0.0054124584		0.0277853124		0.0135149611		0.001307603		0.2455197994		0.2416434597		0.1365479689

		648		0.0052145873		0.0306354538		0.0148159061		0.0013567037		0.2707045492		0.2649039659		0.1416753693

		649		0.0052227611		0.0341744945		0.0168385375		0.0017795478		0.301976631		0.3010680112		0.1858313568

		650		0.0053499243		0.0385811031		0.0188106466		0.0022179724		0.3409148172		0.3363287313		0.2316143625

		651		0.0053474871		0.0434504412		0.021343464		0.0026610855		0.3839418274		0.3816147487		0.2778869594

		652		0.0052836263		0.0486990549		0.0241278913		0.0030371046		0.4303202361		0.4313994747		0.3171531999

		653		0.0054140049		0.0549574383		0.0271717571		0.0038848687		0.4856212892		0.4858228838		0.405681955

		654		0.0052475203		0.0614894144		0.0303864963		0.0041053393		0.5433398944		0.5433014597		0.4287048547

		655		0.005402341		0.0687817931		0.0340711661		0.0053690821		0.6077776567		0.6091822517		0.560672674

		656		0.0055206954		0.076306358		0.037750259		0.0060023875		0.6742670894		0.6749633328		0.6268063312

		657		0.0052399482		0.083548598		0.0415460356		0.0066377744		0.7382618104		0.7428306809		0.6931573521

		658		0.0053845448		0.0910115466		0.0450553521		0.0073214765		0.8042067819		0.8055762112		0.7645537392

		659		0.0054763905		0.0975896567		0.0485645123		0.0081332615		0.8623330417		0.868318944		0.8493253362

		660		0.0052579218		0.1036967337		0.0510311574		0.0085986927		0.9162971037		0.9124218199		0.8979285386

		661		0.0055363472		0.1081333011		0.0532441176		0.0090779178		0.955500015		0.9519888874		0.9479721766

		662		0.0055034449		0.111479871		0.0549660735		0.0091981776		0.985071364		0.9827769422		0.9605304491

		663		0.0052795648		0.1131693348		0.0559293479		0.0095761437		1		1		1

		664		0.0052842028		0.1129908338		0.0555895828		0.0094461506		0.9984227088		0.9939251017		0.9864253213

		665		0.0056410488		0.111578092		0.0548411459		0.0093885874		0.9859392758		0.9805432739		0.9804142184

		666		0.0054255114		0.108061716		0.0531676039		0.0090956548		0.9548674664		0.9506208446		0.9498243876

		667		0.0052308831		0.1034891307		0.0507930145		0.0084453411		0.914462658		0.9081638961		0.8819146211

		668		0.0052096359		0.0977738053		0.0479481109		0.0081179384		0.8639602365		0.8572978709		0.8477252087

		669		0.0053832605		0.0907426029		0.0445824042		0.0072503495		0.8018303112		0.7971200434		0.7571262225

		670		0.00557232		0.0827998668		0.040414393		0.0064513511		0.731645785		0.7225972497		0.6736898804

		671		0.0054786974		0.0744119808		0.036311809		0.0060609207		0.6575277741		0.6492442761		0.6329187247

		672		0.0052274507		0.0665585846		0.0323717929		0.0048694522		0.5881326835		0.5787979687		0.5084982396

		673		0.0056028678		0.0585775711		0.0287717264		0.0045924219		0.5176099264		0.5144298567		0.4795690285

		674		0.0058890185		0.0510437042		0.0249876399		0.0038961261		0.451038298		0.4467715218		0.4068575181

		675		0.0054918313		0.0441859625		0.0214391835		0.0032914428		0.3904411262		0.3833261841		0.3437127576

		676		0.0054794839		0.0376591794		0.0182584133		0.0024987247		0.3327684084		0.326454965		0.260932251

		677		0.0054059317		0.032113947		0.0153196091		0.0017403013		0.2837689824		0.2739100251		0.1817329936

		678		0.0054739001		0.0271605961		0.0131955631		0.0012757982		0.239999609		0.2359327189		0.1332267173

		679		0.0054267435		0.022996299		0.0110744992		0.0009707207		0.2032025638		0.1980087316		0.10136864

		680		0.0058452799		0.019697275		0.0094596958		0.0008337325		0.1740513457		0.1691365292		0.0870634953

		681		0.0057302471		0.0165592562		0.0078128669		0.000458579		0.1463228203		0.1396917225		0.0478876468

		682		0.0057044639		0.0136848865		0.0065068179		0.0002559841		0.1209239812		0.1163399562		0.0267314398

		683		0.0055692522		0.0118703237		0.0056337584		0.0003236175		0.1048899307		0.1007299136		0.0337941338

		684		0.0058448599		0.0099907005		0.0048333546		-9.73E-05		0.0882809865		0.0864189339		-0.0101573845

		685		0.005913212		0.008843259		0.0040129991		-0.0002274781		0.0781418303		0.0717512232		-0.0237546635

		686		0.0055891275		0.0076358183		0.0031817518		-0.0003583728		0.0674725034		0.0568887702		-0.0374234936

		687		0.0056866291		0.0064300178		0.0028176184		-0.0005092924		0.0568176686		0.0503781742		-0.0531834554

		688		0.0058201198		0.0056945495		0.0026926273		-0.0004601122		0.0503188389		0.0481433694		-0.048047758

		689		0.0059322631		0.0053633689		0.0025610058		-0.0005242885		0.0473924222		0.0457900165		-0.0547494344

		690		0.0057926523		0.0048543741		0.001971907		-0.0008198626		0.0428947834		0.0352571066		-0.0856151091

		691		0.005901299		0.0042982087		0.0019222167		-0.0005559335		0.0379803306		0.0343686598		-0.0580540068

		692		0.0059026107		0.0042067948		0.0016774102		-0.0002918439		0.0371725679		0.0299915925		-0.0304761372

		693		0.0058515244		0.0041881329		0.0018605212		-0.000361373		0.0370076659		0.0332655613		-0.0377368018

		694		0.0056446414		0.0036588192		0.001627		-0.0008881707		0.0323304829		0.0290902739		-0.0927482609

		695		0.005611626		0.0031419352		0.0010933229		-0.0008768039		0.0277631323		0.0195482858		-0.0915612765

		696		0.0058149775		0.0031494705		0.0013304779		-0.0006545251		0.0278297169		0.0237885463		-0.0683495451

		697		0.0059100632		0.0036630481		0.0013135488		-0.0003961852		0.0323678508		0.0234858588		-0.0413721002

		698		0.00583085		0.0030450292		0.0011256334		-0.00095404		0.0269068405		0.0201259877		-0.0996267459

		699		0.0059312135		0.0029621928		0.0012714625		-0.0006043274		0.0261748711		0.0227333696		-0.0631075911

		700		0.0059436001		0.0032079071		0.0010578749		-0.0005411454		0.0283460806		0.0189144862		-0.0565097399

		701		0.0061487611		0.0032542502		0.0012540425		-0.000164293		0.0287555826		0.0224219045		-0.0171564932

		702		0.0061611538		0.0028265032		0.0011245173		-0.0003450782		0.0249758758		0.0201060326		-0.0360351894

		703		0.0059697386		0.0029498651		0.0008476084		-0.0005830264		0.0260659398		0.0151549841		-0.0608832177

		704		0.0058907242		0.0029297192		0.0012810426		-0.0003815468		0.0258879242		0.022904658		-0.039843473

		705		0.0057736333		0.0029189561		0.000905762		-0.0008405835		0.0257928184		0.0161947538		-0.0877789172

		706		0.0059722317		0.0032688815		0.0009285379		-0.0005477639		0.0288848695		0.0166019801		-0.0572008847

		707		0.006048216		0.0029759542		0.0010514915		-0.0005180833		0.0262964717		0.0188003529		-0.0541014523

		708		0.0063055321		0.0030867662		0.0008608881		-0.0003480785		0.0272756417		0.0153924216		-0.0363485067

		709		0.0055913562		0.002779268		0.0009108463		-0.000595228		0.0245584899		0.0162856589		-0.0621573847

		710		0.0061876993		0.0029817405		0.0011457475		-0.0004707658		0.0263476013		0.0204856222		-0.0491602729

		711		0.0063944729		0.0030362704		0.0014850515		-0.0002070384		0.0268294445		0.0265522758		-0.0216202262

		712		0.0062225703		0.0026437659		0.0008384789		-0.0004199779		0.0233611507		0.0149917509		-0.0438566839

		713		0.0061043138		0.0028867985		0.0007955569		-0.0002129375		0.0255086634		0.0142243188		-0.022236251

		714		0.0061894585		0.0034748702		0.0007809047		-0.0008616134		0.0307050513		0.0139623427		-0.0899749896

		715		0.0063420413		0.002531192		0.0010081334		-0.0005438342		0.0223664126		0.0180251235		-0.0567905196

		716		0.006474236		0.0030768856		0.0008773067		-0.0003816503		0.0271883333		0.0156859819		-0.0398542773

		717		0.0060514705		0.0026810886		0.0007942083		-0.0003198328		0.0236909463		0.014200207		-0.03339891

		718		0.0065214806		0.0030334292		0.0010330942		-0.0004354425		0.0268043384		0.0184714159		-0.045471591

		719		0.0065735658		0.002763585		0.001207056		5.41E-05		0.0244199099		0.0215817998		0.0056526861

		720		0.0065387976		0.0030639551		0.001069163		3.50E-05		0.0270740755		0.0191163139		0.0036521332

		721		0.0063608224		0.0028101669		0.0005377255		-0.0003723394		0.0248315228		0.0096143702		-0.038881979

		722		0.0066931597		0.0025632191		0.0011065581		0.0002176779		0.0226494134		0.0197849283		0.0227312671

		723		0.0066128573		0.0028570406		0.0012858976		6.21E-05		0.0252457136		0.0229914636		0.0064826708

		724		0.0067295176		0.0027259425		0.0011874582		-0.0001089145		0.0240872895		0.0212313966		-0.0113735233

		725		0.006789411		0.0029421765		0.0012391931		4.06E-05		0.0259980015		0.0221564024		0.0042387725

		726		0.0064367959		0.0029843468		0.0010828128		0.0002738045		0.0263706315		0.0193603685		0.0285923581

		727		0.005698143		0.0021576397		0.0003557671		-0.0002273228		0.0190655862		0.0063610095		-0.0237384539

		728		0.007211132		0.0028728314		0.0011473048		0.0001389781		0.0253852462		0.020513466		0.0145129516

		729		0.006281422		0.0024101096		0.0004331847		-0.0006103245		0.0212964901		0.0077452123		-0.0637338489

		730		0.0064352457		0.0024468654		0.0004323814		-0.0002359124		0.0216212761		0.0077308501		-0.0246354286

		731		0.0067762905		0.0029944079		0.0007776372		-9.37E-05		0.0264595343		0.013903921		-0.0097844289

		732		0.0072501651		0.0029630528		0.000952716		5.80E-05		0.026182471		0.0170342767		0.0060555106

		733		0.0063014873		0.0019616617		0.0005521621		-0.0004978641		0.0173338629		0.0098724924		-0.05199004

		734		0.0070779235		0.0020814748		0.0005949822		-0.0007085419		0.0183925686		0.0106381033		-0.0739903153

		735		0.0070485347		0.0025460075		0.0005887093		6.55E-05		0.0224973268		0.0105259453		0.0068422456

		736		0.006880397		0.0025289007		0.0007369253		-7.78E-05		0.022346166		0.0131760046		-0.008119603

		737		0.0067300699		0.0023044925		8.48E-05		-0.0003339559		0.020363224		0.0015154643		-0.0348737311

		738		0.0065715946		0.0019305891		0.0001588395		-0.0011212259		0.0170592952		0.0028400021		-0.1170853235

		739		0.0069476757		0.0024287996		0.0009035831		-0.0003761672		0.021461641		0.0161557949		-0.0392817039

		740		0.0076891966		0.0029523671		0.0008813272		-0.0001140386		0.0260880483		0.0157578665		-0.0119086142

		741		0.0066509433		0.0019740653		0.000689165		-0.0003852711		0.0174434649		0.0123220644		-0.0402323874

		742		0.0062568928		0.0018876371		0.000176138		-0.0004078231		0.0166797576		0.0031492939		-0.0425874039

		743		0.0068921791		0.0020832175		0.0001509673		-0.0006529742		0.018407968		0.0026992504		-0.0681875982

		744		0.0065091564		0.002220236		0.0006102761		-0.0008149535		0.0196187068		0.0109115545		-0.0851024748

		745		0.0071538645		0.0021520725		0.0004594082		-0.0005503492		0.0190163926		0.0082140813		-0.0574708631

		746		0.006710405		0.0018008359		0.0006959061		-0.0003958748		0.0159127554		0.012442593		-0.0413396932

		747		0.0067334874		0.0018421937		0.0002357557		-0.0005179282		0.0162782052		0.0042152418		-0.0540852533

		748		0.0071191029		0.0016704979		-6.76E-05		-0.0008623885		0.0147610476		-0.0012088267		-0.0900559235

		749		0.0066871662		0.0018098298		0.0001283875		-0.0009795584		0.0159922284		0.0022955299		-0.1022915361

		750		0.0069605908		0.001956461		0.0006032512		-0.0007054923		0.0172879075		0.0107859512		-0.0736718616

		751		0.0070273825		0.0015882349		0.0001328153		-0.0008145918		0.0140341455		0.0023746984		-0.0850647036

		752		0.0074894694		0.001999602		0.0003402749		8.50E-05		0.0176691147		0.0060840131		0.0088807915

		753		0.0064644879		0.0015002729		8.65E-05		-0.0011115693		0.0132568853		0.0015460174		-0.1160769287

		754		0.0069904211		0.0017833686		-0.0005354576		-0.0008849673		0.0157584088		-0.0095738211		-0.0924137482

		755		0.0070873168		0.001239297		-0.0004034266		-0.0003882712		0.0109508197		-0.0072131466		-0.0405456774

		756		0.0067464751		0.0016687829		-0.0001592213		-0.0006932421		0.014745893		-0.0028468291		-0.0723926192

		757		0.0063916096		0.0009305095		-0.0004732482		-0.001465363		0.0082222757		-0.008461536		-0.1530222403

		758		0.0072735134		0.0011379872		0.0004465294		0.0001687575		0.0100556149		0.0079838129		0.0176227021

		759		0.0074483869		0.0011821369		0.000260206		-0.0004542164		0.0104457355		0.0046524059		-0.0474320826

		760		0.007149891		0.0008606028		-0.0007375376		-0.0007403286		0.007604558		-0.0131869521		-0.0773096764

		761		0.0072166063		0.0010392958		-0.0004869009		-0.0003884781		0.0091835461		-0.0087056421		-0.0405672831

		762		0.0072568511		0.0009642607		-0.0010191763		-0.00122511		0.0085205119		-0.0182225677		-0.1279335467

		763		0.0064382409		0.0009065921		0.0001232865		-0.0007752137		0.0080109339		0.0022043253		-0.0809525971

		764		0.0073088668		0.0018207735		5.99E-05		-5.70E-05		0.0160889297		0.0010706074		-0.005957401

		765		0.0072203963		0.0012238252		-0.0003937024		-0.0005653955		0.0108141057		-0.0070392808		-0.0590420877

		766		0.0071475487		0.0013026439		-0.0007416724		-0.0008927689		0.0115105732		-0.0132608797		-0.0932284365

		767		0.0073841107		0.0012023835		8.92E-05		-0.0004895383		0.0106246409		0.0015946246		-0.0511206095

		768		0.0069052242		0.0010458088		-0.0009646815		-0.0011969212		0.0092410968		-0.0172482165		-0.1249898977

		769		0.0072223176		0.0010267891		-0.0005688597		-0.0004608363		0.0090730327		-0.010171042		-0.0481233672

		770		0.0062258793		0.000951341		-0.0004638359		-0.0006722558		0.0084063497		-0.0082932465		-0.0702010935

		771		0.005964437		0.0002266649		-0.0011822582		-0.0007689088		0.0020028828		-0.0211384236		-0.0802941967

		772		0.0068462356		0.001437936		-0.0001560643		-0.0004506996		0.0127060566		-0.0027903837		-0.0470648346

		773		0.0070445095		4.37E-05		-0.0007037866		-0.0007150545		0.0003861268		-0.0125834941		-0.0746704035

		774		0.0072638262		0.0001245294		-0.0007171736		-0.0005460576		0.0011003807		-0.0128228497		-0.0570227024

		775		0.0064320932		4.31E-05		-0.0002446053		-0.001081824		0.0003806366		-0.004373469		-0.1129707309

		776		0.0071122614		0.0013595861		-0.0001126929		-0.0006199405		0.0120137327		-0.002014915		-0.0647380072

		777		0.0075494414		0.0009902827		0.0001169168		-0.0003490097		0.0087504512		0.0020904368		-0.0364457429

		778		0.0069488068		0.001475415		0.0001344725		-0.0002615766		0.0130372328		0.0024043286		-0.0273154406

		779		0.006690531		0.0006660646		-0.000643204		-0.0006915364		0.0058855574		-0.0115002944		-0.0722144976

		780		0.0063049281		0.0003770115		0.0004255667		0.0001857974		0.0033313929		0.0076090045		0.0194021151

		781		0.0074644508		0.0010660749		-2.81E-05		-8.15E-05		0.009420175		-0.0005026208		-0.0085087887

		782		0.0083187725		0.0011221815		0.0011102434		4.06E-05		0.0099159506		0.0198508193		0.0042360691

		783		0.0067111147		7.44E-05		-0.000170762		-0.0005849394		0.0006572165		-0.0030531742		-0.0610829844

		784		0.0063071079		0.0001662974		-0.0003402154		-0.0008197076		0.0014694565		-0.00608295		-0.0855989223

		785		0.0067987451		0.0012925442		-0.0002139725		-0.0005889722		0.0114213288		-0.0038257637		-0.0615041084

		786		0.0076478054		0.0012534194		0.0002136636		-0.0001791457		0.0110756102		0.0038202405		-0.018707501

		787		0.0071111824		0.0011909363		0.0004890282		0.0001527282		0.0105234895		0.0087436773		0.0159488193

		788		0.0066838013		0.000776574		-2.59E-05		-0.0006625893		0.0068620534		-0.0004628196		-0.0691916592

		789		0.0073030749		0.0009160344		-0.0004168746		-0.0006050511		0.0080943693		-0.007453593		-0.06318317

		790		0.0076204068		0.0012237732		0.0001810842		0.0002598952		0.0108136469		0.0032377314		0.0271398629

		791		0.0077395779		0.0009304576		-0.000197879		-0.000793973		0.0082218174		-0.0035380177		-0.0829115631

		792		0.0068267761		-0.0003762707		-0.0014471487		-0.0009316714		-0.0033248468		-0.0258745863		-0.0972908764

		793		0.0073957224		0.0007789079		-0.0006009668		0.0006569389		0.0068826764		-0.0107451065		0.0686016104

		794		0.0076335263		0.0014902204		-0.0002941203		0.0002394853		0.0131680586		-0.0052587832		0.0250085362

		795		0.0081879413		0.0002233757		0.0007871286		0.00080606		0.0019738181		0.0140736234		0.0841737617

		796		0.0067921979		0.0009154637		-0.0002899813		3.55E-05		0.0080893262		-0.0051847796		0.0037116078

		797		0.0073139477		0.0008153964		0.0012606366		0.0012093922		0.0072051005		0.0225398052		0.1262921955

		798		0.0073607308		0.0006295369		5.75E-05		-0.000262094		0.0055627867		0.0010289468		-0.0273694715

		799		0.0070814495		0.0005516437		-0.0003098998		-0.0004577333		0.0048744977		-0.0055409151		-0.0477993306

		800		0.0075571854		0.0008047893		-0.0002264949		-0.0001808535		0.0071113726		-0.0040496615		-0.0188858368

		801		0.0066532562		0.0001499315		-0.0007645676		-0.0007855495		0.0013248423		-0.0136702395		-0.082031927

		802		0.0063765836		0.0005169915		0.0007961015		-0.0009661795		0.0045683006		0.0142340559		-0.1008944273

		803		0.0060943114		0.000118548		-0.0003110379		-0.0003266616		0.0010475278		-0.0055612651		-0.0341120161

		804		0.0060736258		0.0011510941		-0.0009035153		-0.0006177691		0.0101714311		-0.0161545825		-0.0645112585

		805		0.0061042611		0.0005787038		0.0003818305		-0.0004371492		0.0051136098		0.0068270155		-0.0456498188

		806		0.0060338578		-9.86E-05		0.0003161309		-0.0007016157		-0.0008713875		0.0056523264		-0.07326704

		807		0.0062491461		0.0013461353		-0.0007523191		0.000517147		0.0118948767		-0.0134512396		0.0540036812

		808		0.0077264812		0.00085845		0.00102928		0.0010159173		0.0075855356		0.0184032186		0.1060883498

		809		0.0072704069		0.0001201534		-0.000319212		0.0006496281		0.0010617134		-0.0057074146		0.067838169

		810		0.0074872044		0.0005189612		-0.000513119		0.0005276175		0.0045857054		-0.0091744149		0.0550970765

		811		0.0062703392		-0.0004442865		-0.0013486852		-0.0004051852		-0.0039258559		-0.0241140884		-0.0423119368

		812		0.0061637005		-3.62E-06		-0.0009105932		-4.04E-05		-0.000032023		-0.0162811338		-0.0042167531

		813		0.0072767246		0.0004457262		-0.0011562353		-0.0006166318		0.0039385771		-0.0206731419		-0.0643924926

		814		0.0062044254		-0.0012136488		-0.0006892104		-0.0010025961		-0.0107241845		-0.0123228751		-0.1046972707

		815		0.0072027361		0.000302298		0.0005203607		0.0003830482		0.0026712003		0.0093038938		0.0400002593

		816		0.0068770042		0.0004183634		-0.0002064692		-1.24E-05		0.0036967909		-0.0036916069		-0.0012975043

		817		0.0075419345		0.0005322827		0.0001205418		0.0002772236		0.004703418		0.0021552512		0.0289494035

		818		0.0068811332		-0.0003589418		-0.0003165735		-0.0008003293		-0.0031717231		-0.0056602401		-0.0835753186

		819		0.0056334175		-0.0006136332		-0.0001835445		-0.0007229108		-0.0054222568		-0.0032817207		-0.0754908055

		820		0.0060566156		0.000222262		-0.0015643674		-0.0006072225		0.0019639773		-0.0279704216		-0.0634099187

		821		0.0057069291		-0.0002771503		-0.0024162959		-0.0004844186		-0.0024489874		-0.043202648		-0.0505859805

		822		0.0058076582		-0.0014865175		-0.001075575		-0.0018737719		-0.0131353382		-0.0192309593		-0.1956708204

		823		0.0072564823		-0.0005109989		-0.0021017136		0.0001288018		-0.0045153476		-0.0375780094		0.0134502781

		824		0.00757349		0.0012203986		-0.000130601		-0.0001775932		0.0107838275		-0.0023351059		-0.0185453779

		825		0.0060007866		-0.0008392917		-0.0021367471		-0.0017406003		-0.0074162468		-0.038204399		-0.1817642244

		826		0.007528475		-0.0001942049		-0.000140124		0.0003683582		-0.0017160557		-0.0025053757		0.0384662345

		827		0.007321029		0.0005172766		-0.0010343096		-0.0007836374		0.0045708193		-0.0184931465		-0.0818322514

		828		0.0055899927		-0.000546523		-0.0014338874		-0.0008609417		-0.0048292496		-0.0256374776		-0.0899048448

		829		0.0050577382		-0.0018646992		-0.0023125419		-0.002317949		-0.0164770711		-0.0413475577		-0.2420545303

		830		0.006661668		-0.0019281012		-0.0020813618		-0.0009399371		-0.0170373112		-0.0372141266		-0.0981540279

		831		0.0057976102		-0.0013932236		-0.0023553336		-0.002474058		-0.0123109644		-0.0421126607		-0.258356394

		832		0.0063595879		-0.0012799851		-0.0016205407		-0.0009088883		-0.0113103525		-0.0289747835		-0.0949117226

		833		0.005823845		-0.0007023393		-0.0020932639		-0.0012344025		-0.0062060923		-0.037426932		-0.1289039273

		834		0.0055073509		-0.0023986879		-0.0026361314		-0.0031502589		-0.021195564		-0.0471332398		-0.3289694695

		835		0.0067146374		-0.0017122386		-0.001889494		-0.0016873305		-0.0151298817		-0.0337835875		-0.1762014644

		836		0.0072787781		0.0006048065		-0.0019951013		-0.0003897196		0.0053442616		-0.0356718148		-0.0406969203

		837		0.0055343024		-0.0013853795		-0.0039340351		-0.0021221158		-0.0122416507		-0.0703393697		-0.2216044241

		838		0.0054130089		-0.0026353081		-0.0034461063		-0.0027088362		-0.0232864148		-0.0616153498		-0.2828733895

		839		0.0068449993		-8.56E-05		-0.0006993931		0.000609991		-0.0007561287		-0.0125049393		0.0636990197

		840		0.00641943		-0.0003626145		-0.0017669491		-0.0016937252		-0.0032041767		-0.0315925206		-0.1768692378

		841		0.0059720217		-0.0006301248		-0.0038509129		-0.0039248001		-0.005567982		-0.0688531702		-0.4098518367

		842		0.0058392193		-0.0013133816		-0.0007781595		-0.0018928961		-0.0116054548		-0.0139132585		-0.1976678918

		843		0.0056040213		-0.0006867292		-0.0017749411		-0.0013269047		-0.0060681563		-0.0317354157		-0.1385635796

		844		0.0058772112		-7.92E-06		-0.0021475656		-0.0021986156		-0.0000699928		-0.0383978301		-0.2295930018

		845		0.0043324842		-0.0011519496		-0.0036161351		-0.0021319042		-0.0101789909		-0.0646554133		-0.2226265928

		846		0.0062730177		-0.0007499417		-0.0016013012		-0.0020398826		-0.0066267217		-0.0286307871		-0.2130171238

		847		0.0070434571		-0.0016316817		-0.0012736875		-0.000633175		-0.014418055		-0.0227731506		-0.0661200355

		848		0.006261305		-0.000886104		-0.0026596987		-0.0011295913		-0.007829895		-0.047554617		-0.1179588873

		849		0.0059828344		0.0010091712		-0.0011610372		-0.000310624		0.008917356		-0.0207589985		-0.0324372793

		850		0.0049338024		-0.0007747486		-0.002482706		-0.0019179472		-0.0068459236		-0.0443900399		-0.2002838784





Sheet4

		350		350		350		350

		351		351		351		351

		352		352		352		352

		353		353		353		353

		354		354		354		354

		355		355		355		355

		356		356		356		356

		357		357		357		357

		358		358		358		358

		359		359		359		359

		360		360		360		360

		361		361		361		361

		362		362		362		362

		363		363		363		363

		364		364		364		364

		365		365		365		365

		366		366		366		366

		367		367		367		367

		368		368		368		368

		369		369		369		369

		370		370		370		370

		371		371		371		371

		372		372		372		372

		373		373		373		373

		374		374		374		374

		375		375		375		375

		376		376		376		376

		377		377		377		377

		378		378		378		378

		379		379		379		379

		380		380		380		380

		381		381		381		381

		382		382		382		382

		383		383		383		383

		384		384		384		384

		385		385		385		385

		386		386		386		386

		387		387		387		387

		388		388		388		388

		389		389		389		389

		390		390		390		390

		391		391		391		391

		392		392		392		392

		393		393		393		393

		394		394		394		394

		395		395		395		395

		396		396		396		396

		397		397		397		397

		398		398		398		398

		399		399		399		399

		400		400		400		400

		401		401		401		401

		402		402		402		402

		403		403		403		403

		404		404		404		404

		405		405		405		405

		406		406		406		406

		407		407		407		407

		408		408		408		408

		409		409		409		409

		410		410		410		410

		411		411		411		411

		412		412		412		412

		413		413		413		413

		414		414		414		414

		415		415		415		415

		416		416		416		416

		417		417		417		417

		418		418		418		418

		419		419		419		419

		420		420		420		420

		421		421		421		421

		422		422		422		422

		423		423		423		423

		424		424		424		424

		425		425		425		425

		426		426		426		426

		427		427		427		427

		428		428		428		428

		429		429		429		429

		430		430		430		430

		431		431		431		431

		432		432		432		432

		433		433		433		433

		434		434		434		434

		435		435		435		435

		436		436		436		436

		437		437		437		437

		438		438		438		438

		439		439		439		439

		440		440		440		440

		441		441		441		441

		442		442		442		442

		443		443		443		443

		444		444		444		444

		445		445		445		445

		446		446		446		446

		447		447		447		447

		448		448		448		448

		449		449		449		449

		450		450		450		450

		451		451		451		451

		452		452		452		452

		453		453		453		453

		454		454		454		454

		455		455		455		455

		456		456		456		456

		457		457		457		457

		458		458		458		458

		459		459		459		459

		460		460		460		460

		461		461		461		461

		462		462		462		462

		463		463		463		463

		464		464		464		464

		465		465		465		465

		466		466		466		466

		467		467		467		467

		468		468		468		468

		469		469		469		469

		470		470		470		470

		471		471		471		471

		472		472		472		472

		473		473		473		473

		474		474		474		474

		475		475		475		475

		476		476		476		476

		477		477		477		477

		478		478		478		478

		479		479		479		479

		480		480		480		480

		481		481		481		481

		482		482		482		482

		483		483		483		483

		484		484		484		484

		485		485		485		485

		486		486		486		486

		487		487		487		487

		488		488		488		488

		489		489		489		489

		490		490		490		490

		491		491		491		491

		492		492		492		492

		493		493		493		493

		494		494		494		494

		495		495		495		495

		496		496		496		496

		497		497		497		497

		498		498		498		498

		499		499		499		499

		500		500		500		500

		501		501		501		501

		502		502		502		502

		503		503		503		503

		504		504		504		504

		505		505		505		505

		506		506		506		506

		507		507		507		507

		508		508		508		508

		509		509		509		509

		510		510		510		510

		511		511		511		511

		512		512		512		512

		513		513		513		513

		514		514		514		514

		515		515		515		515

		516		516		516		516

		517		517		517		517

		518		518		518		518

		519		519		519		519

		520		520		520		520

		521		521		521		521

		522		522		522		522

		523		523		523		523

		524		524		524		524

		525		525		525		525

		526		526		526		526

		527		527		527		527

		528		528		528		528

		529		529		529		529

		530		530		530		530

		531		531		531		531

		532		532		532		532

		533		533		533		533

		534		534		534		534

		535		535		535		535

		536		536		536		536

		537		537		537		537

		538		538		538		538

		539		539		539		539

		540		540		540		540

		541		541		541		541

		542		542		542		542

		543		543		543		543

		544		544		544		544

		545		545		545		545

		546		546		546		546

		547		547		547		547

		548		548		548		548

		549		549		549		549

		550		550		550		550

		551		551		551		551

		552		552		552		552

		553		553		553		553

		554		554		554		554

		555		555		555		555

		556		556		556		556

		557		557		557		557

		558		558		558		558

		559		559		559		559

		560		560		560		560

		561		561		561		561

		562		562		562		562

		563		563		563		563

		564		564		564		564

		565		565		565		565

		566		566		566		566

		567		567		567		567

		568		568		568		568

		569		569		569		569

		570		570		570		570

		571		571		571		571

		572		572		572		572

		573		573		573		573

		574		574		574		574

		575		575		575		575

		576		576		576		576

		577		577		577		577

		578		578		578		578

		579		579		579		579

		580		580		580		580

		581		581		581		581

		582		582		582		582

		583		583		583		583

		584		584		584		584

		585		585		585		585

		586		586		586		586

		587		587		587		587

		588		588		588		588

		589		589		589		589

		590		590		590		590

		591		591		591		591

		592		592		592		592

		593		593		593		593

		594		594		594		594

		595		595		595		595

		596		596		596		596

		597		597		597		597

		598		598		598		598

		599		599		599		599

		600		600		600		600

		601		601		601		601

		602		602		602		602

		603		603		603		603

		604		604		604		604

		605		605		605		605

		606		606		606		606

		607		607		607		607

		608		608		608		608

		609		609		609		609

		610		610		610		610

		611		611		611		611

		612		612		612		612

		613		613		613		613

		614		614		614		614

		615		615		615		615

		616		616		616		616

		617		617		617		617

		618		618		618		618

		619		619		619		619

		620		620		620		620

		621		621		621		621

		622		622		622		622

		623		623		623		623

		624		624		624		624

		625		625		625		625

		626		626		626		626

		627		627		627		627

		628		628		628		628

		629		629		629		629

		630		630		630		630

		631		631		631		631

		632		632		632		632

		633		633		633		633

		634		634		634		634

		635		635		635		635

		636		636		636		636

		637		637		637		637

		638		638		638		638

		639		639		639		639

		640		640		640		640

		641		641		641		641

		642		642		642		642

		643		643		643		643

		644		644		644		644

		645		645		645		645

		646		646		646		646

		647		647		647		647

		648		648		648		648

		649		649		649		649

		650		650		650		650

		651		651		651		651

		652		652		652		652

		653		653		653		653

		654		654		654		654

		655		655		655		655

		656		656		656		656

		657		657		657		657

		658		658		658		658

		659		659		659		659

		660		660		660		660

		661		661		661		661

		662		662		662		662

		663		663		663		663

		664		664		664		664

		665		665		665		665

		666		666		666		666

		667		667		667		667

		668		668		668		668

		669		669		669		669

		670		670		670		670

		671		671		671		671

		672		672		672		672

		673		673		673		673

		674		674		674		674

		675		675		675		675

		676		676		676		676

		677		677		677		677

		678		678		678		678

		679		679		679		679

		680		680		680		680

		681		681		681		681

		682		682		682		682

		683		683		683		683

		684		684		684		684

		685		685		685		685

		686		686		686		686

		687		687		687		687

		688		688		688		688

		689		689		689		689

		690		690		690		690

		691		691		691		691

		692		692		692		692

		693		693		693		693

		694		694		694		694

		695		695		695		695

		696		696		696		696

		697		697		697		697

		698		698		698		698

		699		699		699		699

		700		700		700		700

		701		701		701		701

		702		702		702		702

		703		703		703		703

		704		704		704		704

		705		705		705		705

		706		706		706		706

		707		707		707		707

		708		708		708		708

		709		709		709		709

		710		710		710		710

		711		711		711		711

		712		712		712		712

		713		713		713		713

		714		714		714		714

		715		715		715		715

		716		716		716		716

		717		717		717		717

		718		718		718		718

		719		719		719		719

		720		720		720		720

		721		721		721		721

		722		722		722		722

		723		723		723		723

		724		724		724		724

		725		725		725		725

		726		726		726		726

		727		727		727		727

		728		728		728		728

		729		729		729		729

		730		730		730		730

		731		731		731		731

		732		732		732		732

		733		733		733		733

		734		734		734		734

		735		735		735		735

		736		736		736		736

		737		737		737		737

		738		738		738		738

		739		739		739		739

		740		740		740		740

		741		741		741		741

		742		742		742		742

		743		743		743		743

		744		744		744		744

		745		745		745		745

		746		746		746		746

		747		747		747		747

		748		748		748		748

		749		749		749		749

		750		750		750		750

		751		751		751		751

		752		752		752		752

		753		753		753		753

		754		754		754		754

		755		755		755		755

		756		756		756		756

		757		757		757		757

		758		758		758		758

		759		759		759		759

		760		760		760		760

		761		761		761		761

		762		762		762		762

		763		763		763		763

		764		764		764		764

		765		765		765		765

		766		766		766		766

		767		767		767		767

		768		768		768		768

		769		769		769		769

		770		770		770		770

		771		771		771		771

		772		772		772		772

		773		773		773		773

		774		774		774		774

		775		775		775		775

		776		776		776		776

		777		777		777		777

		778		778		778		778

		779		779		779		779

		780		780		780		780

		781		781		781		781

		782		782		782		782

		783		783		783		783

		784		784		784		784

		785		785		785		785

		786		786		786		786

		787		787		787		787

		788		788		788		788

		789		789		789		789

		790		790		790		790

		791		791		791		791

		792		792		792		792

		793		793		793		793

		794		794		794		794

		795		795		795		795

		796		796		796		796

		797		797		797		797

		798		798		798		798

		799		799		799		799

		800		800		800		800

		801		801		801		801

		802		802		802		802

		803		803		803		803

		804		804		804		804

		805		805		805		805

		806		806		806		806

		807		807		807		807

		808		808		808		808

		809		809		809		809

		810		810		810		810

		811		811		811		811

		812		812		812		812

		813		813		813		813

		814		814		814		814

		815		815		815		815

		816		816		816		816

		817		817		817		817

		818		818		818		818

		819		819		819		819

		820		820		820		820

		821		821		821		821

		822		822		822		822

		823		823		823		823

		824		824		824		824

		825		825		825		825

		826		826		826		826

		827		827		827		827

		828		828		828		828

		829		829		829		829

		830		830		830		830

		831		831		831		831

		832		832		832		832

		833		833		833		833

		834		834		834		834

		835		835		835		835

		836		836		836		836

		837		837		837		837

		838		838		838		838

		839		839		839		839

		840		840		840		840

		841		841		841		841

		842		842		842		842

		843		843		843		843

		844		844		844		844

		845		845		845		845

		846		846		846		846

		847		847		847		847

		848		848		848		848

		849		849		849		849

		850		850		850		850



zero

Dil1

Dil_2

Dil_3

0.0063626086

0.0331389494

0.0144853266

-0.0032947653

0.0066823554

0.0337790437

0.0152225485

-0.0026904682

0.0063110213

0.0343132541

0.0149697559

-0.0028602274

0.0065864171

0.035688065

0.0159707181

-0.0029013196

0.0064050336

0.0359276235

0.0158831775

-0.0027854387

0.0063150665

0.0367846787

0.016406266

-0.0024634018

0.0064474363

0.0381049477

0.0170311704

-0.0020552382

0.0065289699

0.0392575525

0.0176320933

-0.0021388077

0.0064026695

0.040107809

0.0179326497

-0.0022425002

0.0067532319

0.0414219722

0.0186210591

-0.001773085

0.0069384696

0.0422664955

0.0186691042

-0.0017431269

0.0067374045

0.0430863909

0.0194344483

-0.0019624273

0.0063128602

0.0440889671

0.0197203811

-0.0018561933

0.0066746534

0.0456231758

0.0203016289

-0.0012237678

0.0066737072

0.0462876819

0.0207041018

-0.0014643826

0.0066943429

0.0475868285

0.0215585623

-0.0012944899

0.0072340039

0.0494628772

0.0225698836

-0.0006426353

0.0068470505

0.0507353768

0.0229446366

-0.0008769073

0.006724759

0.0515073948

0.0231299494

-0.0009607556

0.0069450191

0.0537316389

0.0253611654

0.0000983001

0.0067010201

0.0549729765

0.0260424409

0.0001312617

0.007078818

0.0564393178

0.0265987515

0.0007806454

0.0072587198

0.0580190159

0.0277752671

0.0011032363

0.0072689592

0.0592453331

0.0282243136

0.0011183665

0.0069234511

0.0601345152

0.0283672642

0.0011281507

0.0071638124

0.0614098832

0.0290920529

0.0012886496

0.0073622051

0.0620532595

0.0295383912

0.0014517272

0.0067888061

0.0619977787

0.0294227265

0.0011092832

0.0070389584

0.062775895

0.0297508519

0.0012875333

0.0070695304

0.0632137805

0.030065991

0.0013461353

0.0071684439

0.063312836

0.0301994495

0.0014952336

0.0071280757

0.0637229383

0.0304936152

0.0016821137

0.0069729015

0.0634137467

0.0300800297

0.0014257034

0.0068736644

0.0631551295

0.029819468

0.0013722843

0.0072780144

0.0633984134

0.0303414967

0.001633366

0.0071341284

0.06338083

0.0303584021

0.0016921186

0.007377265

0.0632439926

0.0304459948

0.0020730996

0.007266643

0.062912941

0.0300722327

0.0017773125

0.0069699553

0.06231457

0.029481981

0.0008557266

0.0071860766

0.0620795675

0.0293082837

0.0012292766

0.0070838174

0.0621869117

0.0297017321

0.0014750513

0.0069564614

0.061982166

0.0298029706

0.0014048491

0.0071095773

0.0618509911

0.0297805984

0.0015482523

0.0070247515

0.0621226951

0.0298663545

0.0014157566

0.0071298652

0.0627287254

0.0303116292

0.0014904282

0.0068829479

0.0631949753

0.0306862071

0.0016940937

0.0069075911

0.0640979633

0.0308829434

0.001652932

0.0070196479

0.0655749589

0.0317723081

0.0020450358

0.0070487452

0.0671511069

0.0326540656

0.0024403837

0.0071524438

0.0693407655

0.0338404626

0.0028820559

0.0071725757

0.0719268694

0.0351819731

0.0027702542

0.0067546782

0.0740060285

0.035907913

0.002892297

0.0070010489

0.0770745426

0.037602637

0.0033422289

0.007079476

0.0802737251

0.0391416773

0.0037463319

0.0070278826

0.0828142762

0.0403400771

0.0039305002

0.0069124834

0.0856430754

0.0418211706

0.0040787379

0.0070482455

0.0880795717

0.0430644937

0.004435719

0.007000654

0.0901654512

0.0441853926

0.0048211571

0.0067726881

0.0918241218

0.0449572988

0.0048144823

0.0069482545

0.0932991132

0.0456621088

0.0048604212

0.0069893426

0.0943133682

0.0460982099

0.0050634211

0.0069949455

0.0948503017

0.0465275459

0.0051626386

0.0071093403

0.0955780521

0.0467913523

0.0052649179

0.0069454927

0.0954174027

0.0467215851

0.0052294158

0.0072105266

0.0953380465

0.0466528423

0.005409182

0.0067665619

0.0938457474

0.0458790101

0.0052586296

0.0071643386

0.0933086127

0.0462292284

0.0053023356

0.0073758699

0.0935599208

0.0457069427

0.0055166841

0.0068793711

0.0924363136

0.0452758893

0.004907595

0.0067896475

0.092536211

0.0453946739

0.0050778254

0.0068943095

0.0931252912

0.0459887832

0.0052822637

0.0070326966

0.0944798142

0.0464900583

0.0053936658

0.00701094

0.0963877514

0.047556445

0.0056804135

0.0070712143

0.0989458188

0.0489395

0.006053518

0.0070973942

0.101841934

0.0501192287

0.0061670351

0.00717626

0.1057505161

0.0523534678

0.0067870445

0.0070877904

0.1093767658

0.0542823262

0.0071756546

0.0070246202

0.1133921742

0.0562550314

0.0074330084

0.00723074

0.117572464

0.0583253987

0.0080478145

0.0071469699

0.1206941605

0.0597585328

0.0085730823

0.0072196596

0.1236300543

0.0614932626

0.0086252028

0.0071124453

0.1253319681

0.0620385371

0.0089344122

0.0069485963

0.1253379434

0.0620234273

0.0086324112

0.0071322336

0.1241731197

0.061576806

0.0088870367

0.0071783918

0.1212722734

0.0599949248

0.0084856208

0.0070760818

0.1165553257

0.0577188842

0.0079552652

0.0073061027

0.1100220606

0.0543306992

0.0074250819

0.0071101035

0.1022810936

0.0505176634

0.006607811

0.0069205314

0.0930121392

0.0459653027

0.0058149775

0.0072960462

0.0834339634

0.0411435626

0.0051142317

0.0071857609

0.0731842667

0.036229074

0.0039151674

0.006989737

0.0635925606

0.0313099548

0.0030518591

0.0072748293

0.0542573966

0.0271561053

0.0023183634

0.0070522181

0.0457263552

0.0224920157

0.0014584282

0.0073972233

0.0378988124

0.0189895537

0.0007894885

0.007051534

0.0309996642

0.0153506091

0.0000661178

0.0070892903

0.0253759306

0.0128531484

-0.0003709945

0.0071311281

0.0203398764

0.0102354167

-0.0011134801

0.0071483646

0.0162930526

0.0084055671

-0.0014028739

0.0072281607

0.0132118948

0.0068978863

-0.0015737559

0.0068754

0.0102977036

0.0053039077

-0.0019083084

0.0069497013

0.0081343697

0.0043160911

-0.0020531255

0.0070714774

0.0065777446

0.0040526604

-0.0018667611

0.0072846483

0.0054678712

0.0031530689

-0.0019855674

0.0071790498

0.0044341497

0.0027278701

-0.002180123

0.0070928163

0.0034194242

0.0021660686

-0.0023886992

0.0070518763

0.0028439343

0.0018138069

-0.0022584666

0.0073525426

0.0023778595

0.0018010699

-0.0021561689

0.0072014462

0.002052058

0.0015492655

-0.0022530071

0.0070701879

0.0017998742

0.0014181719

-0.0023745398

0.007039879

0.0014818046

0.0012811983

-0.0024465164

0.0072447956

0.0011079595

0.0011576348

-0.0023228412

0.0071706017

0.0011823965

0.0012288613

-0.0022759775

0.0070822388

0.0008601878

0.0010237013

-0.0024898099

0.0073059183

0.0008515248

0.0010849666

-0.0022567154

0.0072318721

0.0008396978

0.0010152427

-0.0024648432

0.0071305232

0.000697021

0.0010218332

-0.002323871

0.0071157082

0.0005045516

0.0009154896

-0.0024182524

0.0070989733

0.0007090516

0.0010465094

-0.0023324706

0.0071602594

0.000817523

0.0009696829

-0.0022125232

0.0071256021

0.0006222265

0.0009489543

-0.0023010063

0.0072764088

0.0006461801

0.0011351063

-0.002279531

0.0071547329

0.0005641108

0.0011687436

-0.0023327796

0.0070751607

0.0004983319

0.0009525862

-0.002278398

0.0072319247

0.000581581

0.0011071292

-0.0021343257

0.0071856817

0.0008054636

0.0011207801

-0.0021261342

0.0071145506

0.0005286802

0.0009957313

-0.0021257736

0.0070083621

0.0004366569

0.000799006

-0.0024751904

0.0072153956

0.0007150151

0.0010744308

-0.0022484234

0.0070241992

0.0007531837

0.0011950895

-0.0020712113

0.0068443683

0.0007591997

0.0011362742

-0.002267634

0.007088948

0.0007354733

0.0011092313

-0.0022894191

0.007093579

0.0010129334

0.0012946211

-0.0019064012

0.0070432206

0.0010002458

0.0012484869

-0.0022573336

0.0070732664

0.0011019647

0.0014226908

-0.0021479262

0.006913404

0.0009797493

0.0012037074

-0.002235444

0.0069320258

0.0009895823

0.0011786328

-0.002187077

0.0070231208

0.0010654521

0.001039685

-0.002580398

0.0066966298

0.0011976074

0.0009519896

-0.0024866182

0.0068685361

0.0013174694

0.0011488361

-0.0023391135

0.0068562021

0.0015166897

0.0012452678

-0.0024997445

0.0070241732

0.001580077

0.0016713035

-0.0020824438

0.0065723308

0.0013866706

0.001305474

-0.0021341711

0.0065321755

0.001693262

0.0015675027

-0.0020455508

0.006686877

0.0017028773

0.0016172305

-0.0021066081

0.006327989

0.0014857528

0.0015733222

-0.0021191791

0.0063599558

0.0017202115

0.0015038574

-0.0019936068

0.0066238176

0.0016785277

0.0018040591

-0.0017928324

0.0068143904

0.0018453392

0.0021333946

-0.0015712283

0.0067692962

0.0017044627

0.0020110444

-0.0017014092

0.0071256286

0.0020498473

0.0022213028

-0.0015096314

0.0069961557

0.0019061485

0.0020069357

-0.0015030276

0.0071245497

0.0019599455

0.0021884425

-0.0012599562

0.0068884711

0.0020699003

0.0020757266

-0.0013339241

0.0069667725

0.0020105764

0.0019324871

-0.0013815605

0.0068237521

0.0019531067

0.0017997182

-0.0014829574

0.0068230154

0.0018937208

0.0016732265

-0.0018002567

0.0067385351

0.0019191747

0.0018197077

-0.001762618

0.0066319918

0.0019441099

0.0016721091

-0.0017346703

0.0069110105

0.0021169283

0.0019232827

-0.0017400846

0.0067201057

0.0017940777

0.0017479943

-0.0017015124

0.0066745221

0.0021174485

0.001853866

-0.0016427191

0.0067631439

0.0021576136

0.0019702429

-0.0014732057

0.0069570662

0.0022462294

0.0019885758

-0.0015481692

0.0065745381

0.0020672993

0.0019529248

-0.0015856202

0.0066250791

0.0024024828

0.0018674365

-0.001745963

0.0067924345

0.0024071161

0.0022187531

-0.0015866519

0.0065992167

0.0026583767

0.0021001506

-0.0017178595

0.0068949931

0.0028487807

0.0022668899

-0.0012942319

0.0068016378

0.0028734046

0.0024825051

-0.0014188711

0.0066106231

0.0030662753

0.0022977004

-0.0014749084

0.0066502336

0.0032033958

0.0023032175

-0.0016084195

0.0066421377

0.0034462723

0.0025365558

-0.0014374478

0.0065820022

0.0036158524

0.0026881993

-0.0012616598

0.0067594368

0.0037394126

0.0026561369

-0.0012630019

0.0066507855

0.0039243358

0.0029007401

-0.0012913413

0.0066321497

0.0039232387

0.0026987484

-0.0013107492

0.006399333

0.0040219082

0.0026400154

-0.0013389823

0.0065650772

0.0040292502

0.0028089425

-0.0011697117

0.0064888713

0.0042096698

0.0029103302

-0.0010517666

0.0063934224

0.0043205097

0.0027884124

-0.0013242724

0.0063913208

0.0043169279

0.0027284173

-0.0011547379

0.0063524432

0.004420002

0.002763116

-0.0012799335

0.0061668251

0.0043653501

0.0028138405

-0.0012258327

0.0060198153

0.0042399652

0.0024772983

-0.0012980516

0.0062376438

0.0043188366

0.0029781435

-0.001049959

0.0060296319

0.0042224256

0.0028343978

-0.0014106146

0.0060608154

0.0044467552

0.0027774705

-0.0011979538

0.0064643831

0.0046199705

0.0027871879

-0.0013959587

0.0062465197

0.0046162815

0.0027679875

-0.0016837205

0.0062750401

0.0046416339

0.0028006313

-0.0014289336

0.0063723279

0.0043640691

0.0027731718

-0.0013211239

0.0063972315

0.0045401291

0.0029093658

-0.0014410599

0.0065267622

0.0047288476

0.0030231064

-0.001326337

0.006621636

0.0046277409

0.0031394321

-0.0012467408

0.0065295738

0.004584129

0.0030801443

-0.0012714161

0.0067503401

0.0048529082

0.0033488027

-0.0010743872

0.0066690543

0.0046317177

0.0033152048

-0.0011092972

0.0064558438

0.0046054507

0.0031194601

-0.0011864404

0.0066660312

0.004762189

0.0035029221

-0.0011025324

0.0066539138

0.004652597

0.0034643023

-0.0009366308

0.0065748538

0.0047849845

0.0034096926

-0.0010014598

0.0065837102

0.0050012516

0.003563703

-0.0008916839

0.006489607

0.0051220111

0.0034459853

-0.0009223206

0.0064041666

0.0050824876

0.0033624722

-0.0009029469

0.0061520431

0.0051316242

0.0034600231

-0.0012671317

0.0064222151

0.0052110241

0.0034612233

-0.0009617887

0.0062488043

0.0054390891

0.0033499766

-0.0010689646

0.0062399283

0.0057413685

0.0034212766

-0.0011359425

0.0061940798

0.0061186738

0.0036605943

-0.0011056308

0.0061704484

0.0062865173

0.0037739836

-0.0010980396

0.0060493182

0.0065231887

0.0039385981

-0.0011241692

0.0061451904

0.0069994177

0.0040238155

-0.001084406

0.0057864608

0.0071755494

0.0044449242

-0.0008808856

0.0061539859

0.0077314614

0.0045368331

-0.0008667287

0.005997532

0.0080164615

0.0047593098

-0.0009751161

0.0060569569

0.0084756697

0.0048144301

-0.0007915958

0.006177275

0.0089301048

0.0051906947

-0.0006285222

0.0058336048

0.0090569044

0.0053759487

-0.0007559885

0.0058061364

0.0100325812

0.0053023878

-0.0007691155

0.0056898552

0.0101661421

0.0056946808

-0.0008533979

0.0057484247

0.0103612458

0.0056425696

-0.0007588827

0.0059117163

0.0108735273

0.0058272029

-0.0006724625

0.0057182079

0.0113429688

0.0062479377

-0.0006077912

0.0057008183

0.0113076838

0.0060650678

-0.0005764602

0.0058653252

0.0116846468

0.0062050293

-0.0006136849

0.0058822227

0.0119754728

0.0063754539

-0.0006444447

0.0057519921

0.0121139716

0.0062858341

-0.0006419633

0.0059043425

0.0122988215

0.0064939423

-0.0007442049

0.0058376449

0.0125236586

0.006677282

-0.000635036

0.0060197893

0.0128283789

0.0066223717

-0.0005645683

0.0060512871

0.0129344529

0.0067774998

-0.000430167

0.0059544384

0.0129595222

0.0067801815

-0.0006246449

0.0060747806

0.0131311249

0.006619297

-0.0005377844

0.0058824853

0.0129805384

0.0064847721

-0.0005511765

0.0059282482

0.0128837861

0.0066569103

-0.0005408351

0.005922344

0.0128368577

0.0065073171

-0.0006092905

0.0059157573

0.0125934435

0.0062200758

-0.0006737548

0.0060184244

0.0125704473

0.0063423039

-0.0005067586

0.0056754826

0.0121107772

0.0059729666

-0.0007705109

0.0055845389

0.0119466018

0.0058480087

-0.0007321625

0.0058773947

0.0119600929

0.0061975196

-0.000521341

0.0056746174

0.0117715839

0.0058348901

-0.0006005532

0.0056793904

0.011812998

0.0060518906

-0.0005219098

0.0058547254

0.0119768297

0.0057496573

-0.0004263914

0.00567029

0.0120078046

0.0059008789

-0.0003486993

0.0056547644

0.0121917035

0.0060811071

-0.0005812168

0.0056610848

0.0126121771

0.0063042189

-0.0003120726

0.0054750275

0.0129445605

0.0061425911

-0.0004811606

0.0054219468

0.0136858746

0.00659806

-0.0004003749

0.0053899176

0.0142123401

0.0068754791

-0.0004859183

0.0054444103

0.0146625666

0.0071769971

-0.0002705794

0.0054828133

0.0154321948

0.0072237127

-0.0003783915

0.0055087404

0.0161275845

0.0077548618

-0.0001901167

0.0054706237

0.017012354

0.0083096428

0.0001300188

0.0055588693

0.0176915675

0.0085958941

0.0000536649

0.0055543333

0.0182017218

0.0088660847

0.0000910246

0.0055473591

0.0190117024

0.0093152169

0.0001349904

0.0055509508

0.0198028199

0.0097113652

0.0002464526

0.0054692342

0.0202619471

0.0097216889

0.0001160623

0.005637587

0.0210983623

0.0103522846

0.0003749389

0.0053571574

0.0213187765

0.0102322362

0.0004090096

0.0052323765

0.02175273

0.0104153054

0.0000665061

0.0054254327

0.022242859

0.0107047521

0.0002989046

0.0054379883

0.0223962553

0.0110956896

0.0003890073

0.0053262864

0.0226779822

0.0108439866

0.0004300493

0.0055094482

0.0229166914

0.0108586643

0.0004391703

0.0055114147

0.0229393411

0.010971427

0.0004789992

0.0052438783

0.0229420699

0.0109788598

0.0004056672

0.0051972703

0.0227171481

0.0110963536

0.0004398959

0.0054638078

0.0227173939

0.0108214812

0.0005245593

0.0054097064

0.0227276497

0.0109900376

0.0004503645

0.0054363371

0.0222447384

0.0107454537

0.0004616108

0.0054697846

0.0216616206

0.0102695273

0.0002099342

0.0053207045

0.0213188306

0.0101611074

0.0003635653

0.0052827089

0.0209740829

0.009910184

0.0003107687

0.0052697915

0.020278953

0.0096662864

0.0001674109

0.0053277276

0.0197816864

0.0093487799

0.0002164865

0.0052110241

0.0191721581

0.0091549233

0.0001999634

0.00525842

0.0183744635

0.0088097611

-0.0000037276

0.0052906484

0.0180197414

0.0083946148

-0.0000430721

0.0055256765

0.0175960772

0.0086155646

0.0000906103

0.0053404374

0.0167816021

0.0080025652

0.0001270669

0.005270394

0.0163939819

0.0078695361

-0.0003290413

0.0051209894

0.016154021

0.0075034015

-0.0001936874

0.0053580222

0.0159642994

0.0075504156

-0.0003171943

0.0053817406

0.0159255508

0.0074302694

-0.0000563242

0.0053559258

0.0158863738

0.0075308718

-0.0001414697

0.005306161

0.0160572566

0.007682241

-0.0000635713

0.005404071

0.016342802

0.0078550912

0.000057626

0.0053970469

0.016942367

0.0079654418

0.0000653152

0.0052469959

0.0174250193

0.0083763255

-0.0000909538

0.0054860641

0.0187848043

0.0088258497

0.0002911076

0.0056288545

0.0199712124

0.0095744235

0.0005596267

0.0051608048

0.0210323054

0.0100498265

0.0004522562

0.005377023

0.0228840001

0.0111296298

0.0004928377

0.0052400269

0.0251631848

0.0120711448

0.0006934171

0.0054124584

0.0277853124

0.0135149611

0.001307603

0.0052145873

0.0306354538

0.0148159061

0.0013567037

0.0052227611

0.0341744945

0.0168385375

0.0017795478

0.0053499243

0.0385811031

0.0188106466

0.0022179724

0.0053474871

0.0434504412

0.021343464

0.0026610855

0.0052836263

0.0486990549

0.0241278913

0.0030371046

0.0054140049

0.0549574383

0.0271717571

0.0038848687

0.0052475203

0.0614894144

0.0303864963

0.0041053393

0.005402341

0.0687817931

0.0340711661

0.0053690821

0.0055206954

0.076306358

0.037750259

0.0060023875

0.0052399482

0.083548598

0.0415460356

0.0066377744

0.0053845448

0.0910115466

0.0450553521

0.0073214765

0.0054763905

0.0975896567

0.0485645123

0.0081332615

0.0052579218

0.1036967337

0.0510311574

0.0085986927

0.0055363472

0.1081333011

0.0532441176

0.0090779178

0.0055034449

0.111479871

0.0549660735

0.0091981776

0.0052795648

0.1131693348

0.0559293479

0.0095761437

0.0052842028

0.1129908338

0.0555895828

0.0094461506

0.0056410488

0.111578092

0.0548411459

0.0093885874

0.0054255114

0.108061716

0.0531676039

0.0090956548

0.0052308831

0.1034891307

0.0507930145

0.0084453411

0.0052096359

0.0977738053

0.0479481109

0.0081179384

0.0053832605

0.0907426029

0.0445824042

0.0072503495

0.00557232

0.0827998668

0.040414393

0.0064513511

0.0054786974

0.0744119808

0.036311809

0.0060609207

0.0052274507

0.0665585846

0.0323717929

0.0048694522

0.0056028678

0.0585775711

0.0287717264

0.0045924219

0.0058890185

0.0510437042

0.0249876399

0.0038961261

0.0054918313

0.0441859625

0.0214391835

0.0032914428

0.0054794839

0.0376591794

0.0182584133

0.0024987247

0.0054059317

0.032113947

0.0153196091

0.0017403013

0.0054739001

0.0271605961

0.0131955631

0.0012757982

0.0054267435

0.022996299

0.0110744992

0.0009707207

0.0058452799

0.019697275

0.0094596958

0.0008337325

0.0057302471

0.0165592562

0.0078128669

0.000458579

0.0057044639

0.0136848865

0.0065068179

0.0002559841

0.0055692522

0.0118703237

0.0056337584

0.0003236175

0.0058448599

0.0099907005

0.0048333546

-0.0000972686

0.005913212

0.008843259

0.0040129991

-0.0002274781

0.0055891275

0.0076358183

0.0031817518

-0.0003583728

0.0056866291

0.0064300178

0.0028176184

-0.0005092924

0.0058201198

0.0056945495

0.0026926273

-0.0004601122

0.0059322631

0.0053633689

0.0025610058

-0.0005242885

0.0057926523

0.0048543741

0.001971907

-0.0008198626

0.005901299

0.0042982087

0.0019222167

-0.0005559335

0.0059026107

0.0042067948

0.0016774102

-0.0002918439

0.0058515244

0.0041881329

0.0018605212

-0.000361373

0.0056446414

0.0036588192

0.001627

-0.0008881707

0.005611626

0.0031419352

0.0010933229

-0.0008768039

0.0058149775

0.0031494705

0.0013304779

-0.0006545251

0.0059100632

0.0036630481

0.0013135488

-0.0003961852

0.00583085

0.0030450292

0.0011256334

-0.00095404

0.0059312135

0.0029621928

0.0012714625

-0.0006043274

0.0059436001

0.0032079071

0.0010578749

-0.0005411454

0.0061487611

0.0032542502

0.0012540425

-0.000164293

0.0061611538

0.0028265032

0.0011245173

-0.0003450782

0.0059697386

0.0029498651

0.0008476084

-0.0005830264

0.0058907242

0.0029297192

0.0012810426

-0.0003815468

0.0057736333

0.0029189561

0.000905762

-0.0008405835

0.0059722317

0.0032688815

0.0009285379

-0.0005477639

0.006048216

0.0029759542

0.0010514915

-0.0005180833

0.0063055321

0.0030867662

0.0008608881

-0.0003480785

0.0055913562

0.002779268

0.0009108463

-0.000595228

0.0061876993

0.0029817405

0.0011457475

-0.0004707658

0.0063944729

0.0030362704

0.0014850515

-0.0002070384

0.0062225703

0.0026437659

0.0008384789

-0.0004199779

0.0061043138

0.0028867985

0.0007955569

-0.0002129375

0.0061894585

0.0034748702

0.0007809047

-0.0008616134

0.0063420413

0.002531192

0.0010081334

-0.0005438342

0.006474236

0.0030768856

0.0008773067

-0.0003816503

0.0060514705

0.0026810886

0.0007942083

-0.0003198328

0.0065214806

0.0030334292

0.0010330942

-0.0004354425

0.0065735658

0.002763585

0.001207056

0.0000541309

0.0065387976

0.0030639551

0.001069163

0.0000349734

0.0063608224

0.0028101669

0.0005377255

-0.0003723394

0.0066931597

0.0025632191

0.0011065581

0.0002176779

0.0066128573

0.0028570406

0.0012858976

0.000062079

0.0067295176

0.0027259425

0.0011874582

-0.0001089145

0.006789411

0.0029421765

0.0012391931

0.0000405911

0.0064367959

0.0029843468

0.0010828128

0.0002738045

0.005698143

0.0021576397

0.0003557671

-0.0002273228

0.007211132

0.0028728314

0.0011473048

0.0001389781

0.006281422

0.0024101096

0.0004331847

-0.0006103245

0.0064352457

0.0024468654

0.0004323814

-0.0002359124

0.0067762905

0.0029944079

0.0007776372

-0.0000936971

0.0072501651

0.0029630528

0.000952716

0.0000579884

0.0063014873

0.0019616617

0.0005521621

-0.0004978641

0.0070779235

0.0020814748

0.0005949822

-0.0007085419

0.0070485347

0.0025460075

0.0005887093

0.0000655223

0.006880397

0.0025289007

0.0007369253

-0.0000777545

0.0067300699

0.0023044925

0.0000847589

-0.0003339559

0.0065715946

0.0019305891

0.0001588395

-0.0011212259

0.0069476757

0.0024287996

0.0009035831

-0.0003761672

0.0076891966

0.0029523671

0.0008813272

-0.0001140386

0.0066509433

0.0019740653

0.000689165

-0.0003852711

0.0062568928

0.0018876371

0.000176138

-0.0004078231

0.0068921791

0.0020832175

0.0001509673

-0.0006529742

0.0065091564

0.002220236

0.0006102761

-0.0008149535

0.0071538645

0.0021520725

0.0004594082

-0.0005503492

0.006710405

0.0018008359

0.0006959061

-0.0003958748

0.0067334874

0.0018421937

0.0002357557

-0.0005179282

0.0071191029

0.0016704979

-0.0000676089

-0.0008623885

0.0066871662

0.0018098298

0.0001283875

-0.0009795584

0.0069605908

0.001956461

0.0006032512

-0.0007054923

0.0070273825

0.0015882349

0.0001328153

-0.0008145918

0.0074894694

0.001999602

0.0003402749

0.0000850437

0.0064644879

0.0015002729

0.0000864677

-0.0011115693

0.0069904211

0.0017833686

-0.0005354576

-0.0008849673

0.0070873168

0.001239297

-0.0004034266

-0.0003882712

0.0067464751

0.0016687829

-0.0001592213

-0.0006932421

0.0063916096

0.0009305095

-0.0004732482

-0.001465363

0.0072735134

0.0011379872

0.0004465294

0.0001687575

0.0074483869

0.0011821369

0.000260206

-0.0004542164

0.007149891

0.0008606028

-0.0007375376

-0.0007403286

0.0072166063

0.0010392958

-0.0004869009

-0.0003884781

0.0072568511

0.0009642607

-0.0010191763

-0.00122511

0.0064382409

0.0009065921

0.0001232865

-0.0007752137

0.0073088668

0.0018207735

0.0000598784

-0.0000570489

0.0072203963

0.0012238252

-0.0003937024

-0.0005653955

0.0071475487

0.0013026439

-0.0007416724

-0.0008927689

0.0073841107

0.0012023835

0.0000891863

-0.0004895383

0.0069052242

0.0010458088

-0.0009646815

-0.0011969212

0.0072223176

0.0010267891

-0.0005688597

-0.0004608363

0.0062258793

0.000951341

-0.0004638359

-0.0006722558

0.005964437

0.0002266649

-0.0011822582

-0.0007689088

0.0068462356

0.001437936

-0.0001560643

-0.0004506996

0.0070445095

0.0000436977

-0.0007037866

-0.0007150545

0.0072638262

0.0001245294

-0.0007171736

-0.0005460576

0.0064320932

0.0000430764

-0.0002446053

-0.001081824

0.0071122614

0.0013595861

-0.0001126929

-0.0006199405

0.0075494414

0.0009902827

0.0001169168

-0.0003490097

0.0069488068

0.001475415

0.0001344725

-0.0002615766

0.006690531

0.0006660646

-0.000643204

-0.0006915364

0.0063049281

0.0003770115

0.0004255667

0.0001857974

0.0074644508

0.0010660749

-0.0000281113

-0.0000814814

0.0083187725

0.0011221815

0.0011102434

0.0000405652

0.0067111147

0.0000743768

-0.000170762

-0.0005849394

0.0063071079

0.0001662974

-0.0003402154

-0.0008197076

0.0067987451

0.0012925442

-0.0002139725

-0.0005889722

0.0076478054

0.0012534194

0.0002136636

-0.0001791457

0.0071111824

0.0011909363

0.0004890282

0.0001527282

0.0066838013

0.000776574

-0.0000258852

-0.0006625893

0.0073030749

0.0009160344

-0.0004168746

-0.0006050511

0.0076204068

0.0012237732

0.0001810842

0.0002598952

0.0077395779

0.0009304576

-0.000197879

-0.000793973

0.0068267761

-0.0003762707

-0.0014471487

-0.0009316714

0.0073957224

0.0007789079

-0.0006009668

0.0006569389

0.0076335263

0.0014902204

-0.0002941203

0.0002394853

0.0081879413

0.0002233757

0.0007871286

0.00080606

0.0067921979

0.0009154637

-0.0002899813

0.0000355429

0.0073139477

0.0008153964

0.0012606366

0.0012093922

0.0073607308

0.0006295369

0.0000575483

-0.000262094

0.0070814495

0.0005516437

-0.0003098998

-0.0004577333

0.0075571854

0.0008047893

-0.0002264949

-0.0001808535

0.0066532562

0.0001499315

-0.0007645676

-0.0007855495

0.0063765836

0.0005169915

0.0007961015

-0.0009661795

0.0060943114

0.000118548

-0.0003110379

-0.0003266616

0.0060736258

0.0011510941

-0.0009035153

-0.0006177691

0.0061042611

0.0005787038

0.0003818305

-0.0004371492

0.0060338578

-0.0000986143

0.0003161309

-0.0007016157

0.0062491461

0.0013461353

-0.0007523191

0.000517147

0.0077264812

0.00085845

0.00102928

0.0010159173

0.0072704069

0.0001201534

-0.000319212

0.0006496281

0.0074872044

0.0005189612

-0.000513119

0.0005276175

0.0062703392

-0.0004442865

-0.0013486852

-0.0004051852

0.0061637005

-0.000003624

-0.0009105932

-0.0000403802

0.0072767246

0.0004457262

-0.0011562353

-0.0006166318

0.0062044254

-0.0012136488

-0.0006892104

-0.0010025961

0.0072027361

0.000302298

0.0005203607

0.0003830482

0.0068770042

0.0004183634

-0.0002064692

-0.0000124251

0.0075419345

0.0005322827

0.0001205418

0.0002772236

0.0068811332

-0.0003589418

-0.0003165735

-0.0008003293

0.0056334175

-0.0006136332

-0.0001835445

-0.0007229108

0.0060566156

0.000222262

-0.0015643674

-0.0006072225

0.0057069291

-0.0002771503

-0.0024162959

-0.0004844186

0.0058076582

-0.0014865175

-0.001075575

-0.0018737719

0.0072564823

-0.0005109989

-0.0021017136

0.0001288018

0.00757349

0.0012203986

-0.000130601

-0.0001775932

0.0060007866

-0.0008392917

-0.0021367471

-0.0017406003

0.007528475

-0.0001942049

-0.000140124

0.0003683582

0.007321029

0.0005172766

-0.0010343096

-0.0007836374

0.0055899927

-0.000546523

-0.0014338874

-0.0008609417

0.0050577382

-0.0018646992

-0.0023125419

-0.002317949

0.006661668

-0.0019281012

-0.0020813618

-0.0009399371

0.0057976102

-0.0013932236

-0.0023553336

-0.002474058

0.0063595879

-0.0012799851

-0.0016205407

-0.0009088883

0.005823845

-0.0007023393

-0.0020932639

-0.0012344025

0.0055073509

-0.0023986879

-0.0026361314

-0.0031502589

0.0067146374

-0.0017122386

-0.001889494

-0.0016873305

0.0072787781

0.0006048065

-0.0019951013

-0.0003897196

0.0055343024

-0.0013853795

-0.0039340351

-0.0021221158

0.0054130089

-0.0026353081

-0.0034461063

-0.0027088362

0.0068449993

-0.0000855706

-0.0006993931

0.000609991

0.00641943

-0.0003626145

-0.0017669491

-0.0016937252

0.0059720217

-0.0006301248

-0.0038509129

-0.0039248001

0.0058392193

-0.0013133816

-0.0007781595

-0.0018928961

0.0056040213

-0.0006867292

-0.0017749411

-0.0013269047

0.0058772112

-0.000007921

-0.0021475656

-0.0021986156

0.0043324842

-0.0011519496

-0.0036161351

-0.0021319042

0.0062730177

-0.0007499417

-0.0016013012

-0.0020398826

0.0070434571

-0.0016316817

-0.0012736875

-0.000633175

0.006261305

-0.000886104

-0.0026596987

-0.0011295913

0.0059828344

0.0010091712

-0.0011610372

-0.000310624

0.0049338024

-0.0007747486

-0.002482706

-0.0019179472



Sheet5

		350		350		350

		351		351		351

		352		352		352

		353		353		353

		354		354		354

		355		355		355

		356		356		356

		357		357		357

		358		358		358

		359		359		359

		360		360		360

		361		361		361

		362		362		362

		363		363		363

		364		364		364

		365		365		365

		366		366		366

		367		367		367

		368		368		368

		369		369		369

		370		370		370

		371		371		371

		372		372		372

		373		373		373

		374		374		374

		375		375		375

		376		376		376

		377		377		377

		378		378		378

		379		379		379

		380		380		380

		381		381		381

		382		382		382

		383		383		383

		384		384		384

		385		385		385

		386		386		386

		387		387		387

		388		388		388

		389		389		389

		390		390		390

		391		391		391

		392		392		392

		393		393		393

		394		394		394

		395		395		395

		396		396		396

		397		397		397

		398		398		398

		399		399		399

		400		400		400

		401		401		401

		402		402		402

		403		403		403

		404		404		404

		405		405		405

		406		406		406

		407		407		407

		408		408		408

		409		409		409

		410		410		410

		411		411		411

		412		412		412

		413		413		413

		414		414		414

		415		415		415

		416		416		416

		417		417		417

		418		418		418

		419		419		419

		420		420		420

		421		421		421

		422		422		422

		423		423		423

		424		424		424

		425		425		425

		426		426		426

		427		427		427

		428		428		428

		429		429		429

		430		430		430

		431		431		431

		432		432		432

		433		433		433

		434		434		434

		435		435		435

		436		436		436

		437		437		437

		438		438		438

		439		439		439

		440		440		440

		441		441		441

		442		442		442

		443		443		443

		444		444		444

		445		445		445

		446		446		446

		447		447		447

		448		448		448

		449		449		449

		450		450		450

		451		451		451

		452		452		452

		453		453		453

		454		454		454

		455		455		455

		456		456		456

		457		457		457

		458		458		458

		459		459		459

		460		460		460

		461		461		461

		462		462		462

		463		463		463

		464		464		464

		465		465		465

		466		466		466

		467		467		467

		468		468		468

		469		469		469

		470		470		470

		471		471		471

		472		472		472

		473		473		473

		474		474		474

		475		475		475

		476		476		476

		477		477		477

		478		478		478

		479		479		479

		480		480		480

		481		481		481

		482		482		482

		483		483		483

		484		484		484

		485		485		485

		486		486		486

		487		487		487

		488		488		488

		489		489		489

		490		490		490

		491		491		491

		492		492		492

		493		493		493

		494		494		494

		495		495		495

		496		496		496

		497		497		497

		498		498		498

		499		499		499

		500		500		500

		501		501		501

		502		502		502

		503		503		503

		504		504		504

		505		505		505

		506		506		506

		507		507		507

		508		508		508

		509		509		509

		510		510		510

		511		511		511

		512		512		512

		513		513		513

		514		514		514

		515		515		515

		516		516		516

		517		517		517

		518		518		518

		519		519		519

		520		520		520

		521		521		521

		522		522		522

		523		523		523

		524		524		524

		525		525		525

		526		526		526

		527		527		527

		528		528		528

		529		529		529

		530		530		530

		531		531		531

		532		532		532

		533		533		533

		534		534		534

		535		535		535

		536		536		536

		537		537		537

		538		538		538

		539		539		539

		540		540		540

		541		541		541

		542		542		542

		543		543		543

		544		544		544

		545		545		545

		546		546		546

		547		547		547

		548		548		548

		549		549		549

		550		550		550

		551		551		551

		552		552		552

		553		553		553

		554		554		554

		555		555		555

		556		556		556

		557		557		557

		558		558		558

		559		559		559

		560		560		560

		561		561		561

		562		562		562

		563		563		563

		564		564		564

		565		565		565

		566		566		566

		567		567		567

		568		568		568

		569		569		569

		570		570		570

		571		571		571

		572		572		572

		573		573		573

		574		574		574

		575		575		575

		576		576		576

		577		577		577

		578		578		578

		579		579		579

		580		580		580

		581		581		581

		582		582		582

		583		583		583

		584		584		584

		585		585		585

		586		586		586

		587		587		587

		588		588		588

		589		589		589

		590		590		590

		591		591		591

		592		592		592

		593		593		593

		594		594		594

		595		595		595

		596		596		596

		597		597		597

		598		598		598

		599		599		599

		600		600		600

		601		601		601

		602		602		602

		603		603		603

		604		604		604

		605		605		605

		606		606		606

		607		607		607

		608		608		608

		609		609		609

		610		610		610

		611		611		611

		612		612		612

		613		613		613

		614		614		614

		615		615		615

		616		616		616

		617		617		617

		618		618		618

		619		619		619

		620		620		620

		621		621		621

		622		622		622

		623		623		623

		624		624		624

		625		625		625

		626		626		626

		627		627		627

		628		628		628

		629		629		629

		630		630		630

		631		631		631

		632		632		632

		633		633		633

		634		634		634

		635		635		635

		636		636		636

		637		637		637

		638		638		638

		639		639		639

		640		640		640

		641		641		641

		642		642		642

		643		643		643

		644		644		644

		645		645		645

		646		646		646

		647		647		647

		648		648		648

		649		649		649

		650		650		650

		651		651		651

		652		652		652

		653		653		653

		654		654		654

		655		655		655

		656		656		656

		657		657		657

		658		658		658

		659		659		659

		660		660		660

		661		661		661

		662		662		662

		663		663		663

		664		664		664

		665		665		665

		666		666		666

		667		667		667

		668		668		668

		669		669		669

		670		670		670

		671		671		671

		672		672		672

		673		673		673

		674		674		674

		675		675		675

		676		676		676

		677		677		677

		678		678		678

		679		679		679

		680		680		680

		681		681		681

		682		682		682

		683		683		683

		684		684		684

		685		685		685

		686		686		686

		687		687		687

		688		688		688

		689		689		689

		690		690		690

		691		691		691

		692		692		692

		693		693		693

		694		694		694

		695		695		695

		696		696		696

		697		697		697

		698		698		698

		699		699		699

		700		700		700

		701		701		701

		702		702		702

		703		703		703

		704		704		704

		705		705		705

		706		706		706

		707		707		707

		708		708		708

		709		709		709

		710		710		710

		711		711		711

		712		712		712

		713		713		713

		714		714		714

		715		715		715

		716		716		716

		717		717		717

		718		718		718

		719		719		719

		720		720		720

		721		721		721

		722		722		722

		723		723		723

		724		724		724

		725		725		725

		726		726		726

		727		727		727

		728		728		728

		729		729		729

		730		730		730

		731		731		731

		732		732		732

		733		733		733

		734		734		734

		735		735		735

		736		736		736

		737		737		737

		738		738		738

		739		739		739

		740		740		740

		741		741		741

		742		742		742

		743		743		743

		744		744		744

		745		745		745

		746		746		746

		747		747		747

		748		748		748

		749		749		749

		750		750		750

		751		751		751

		752		752		752

		753		753		753

		754		754		754

		755		755		755

		756		756		756

		757		757		757

		758		758		758

		759		759		759

		760		760		760

		761		761		761

		762		762		762

		763		763		763

		764		764		764

		765		765		765

		766		766		766

		767		767		767

		768		768		768

		769		769		769

		770		770		770

		771		771		771

		772		772		772

		773		773		773

		774		774		774

		775		775		775

		776		776		776

		777		777		777

		778		778		778

		779		779		779

		780		780		780

		781		781		781

		782		782		782

		783		783		783

		784		784		784

		785		785		785

		786		786		786

		787		787		787

		788		788		788

		789		789		789

		790		790		790

		791		791		791

		792		792		792

		793		793		793

		794		794		794

		795		795		795

		796		796		796

		797		797		797

		798		798		798

		799		799		799

		800		800		800

		801		801		801

		802		802		802

		803		803		803

		804		804		804

		805		805		805

		806		806		806

		807		807		807

		808		808		808

		809		809		809

		810		810		810

		811		811		811

		812		812		812

		813		813		813

		814		814		814

		815		815		815

		816		816		816

		817		817		817

		818		818		818

		819		819		819

		820		820		820

		821		821		821

		822		822		822

		823		823		823

		824		824		824

		825		825		825

		826		826		826

		827		827		827

		828		828		828

		829		829		829

		830		830		830

		831		831		831

		832		832		832

		833		833		833

		834		834		834

		835		835		835

		836		836		836

		837		837		837

		838		838		838

		839		839		839

		840		840		840

		841		841		841

		842		842		842

		843		843		843

		844		844		844

		845		845		845

		846		846		846

		847		847		847

		848		848		848

		849		849		849

		850		850		850



Dil1

Dil_2

Dil_3

0.2928262279

0.2589933041

-0.3440597129

0.2984823032

0.2721746112

-0.2809552886

0.3032027547

0.267654755

-0.2986825864

0.3153510185

0.2855516599

-0.3029736958

0.3174678333

0.2839864601

-0.2908726895

0.3250410437

0.293339126

-0.2572436086

0.3367073574

0.3045122298

-0.2146206551

0.3468921377

0.3152565505

-0.2233474913

0.3544052731

0.3206304093

-0.2341757043

0.3660176345

0.3329389639

-0.1851564718

0.3734801093

0.3337979953

-0.1820280635

0.3807249642

0.3474821193

-0.2049287631

0.3895840443

0.3525945122

-0.1938351548

0.4031407968

0.3629870476

-0.1277933787

0.4090125827

0.3701831428

-0.1529198679

0.4204922522

0.3854606412

-0.1351786263

0.4370696116

0.4035427637

-0.0671079401

0.4483138199

0.4102432357

-0.0915720656

0.4551356147

0.4135565715

-0.1003280235

0.4747897391

0.4534500466

0.0102651075

0.4857585895

0.4656310488

0.013707156

0.4987156447

0.4755777164

0.0815198094

0.5126743569

0.4966134632

0.1152067445

0.5235104826

0.50464228

0.1167867281

0.5313675765

0.5071981934

0.117808459

0.5426371314

0.5201571994

0.1345687436

0.5483222073

0.528137593

0.1515983137

0.5478319619

0.5260695431

0.1158381933

0.5547076433

0.5319363276

0.1344521721

0.5585769381

0.5375709198

0.1405717501

0.5594522235

0.5399571184

0.1561415188

0.5630760171

0.5452167129

0.1756566937

0.560343902

0.5378219287

0.1488807447

0.5580586796

0.5331631634

0.1433023876

0.5602084122

0.5424968795

0.1705661527

0.56005304

0.5427991424

0.1767014617

0.5588439019

0.5443652746

0.2164858496

0.5559186249

0.5376825201

0.185597934

0.5506312294

0.5271289953

0.0893602503

0.5485546731

0.5240233404

0.128368651

0.5495032008

0.5310580801

0.154033954

0.5476940032

0.532868194

0.146703011

0.5465349002

0.5324681852

0.1616780564

0.548935763

0.53400148

0.1478420429

0.5542908375

0.5419628573

0.1556397101

0.5584107689

0.5486601988

0.1769077141

0.5663898563

0.5521777849

0.1726093592

0.5794410565

0.5680793583

0.213555257

0.593368398

0.5838449201

0.2548399163

0.6127169131

0.6050573421

0.3009620604

0.6355685449

0.6290431494

0.2892870286

0.653940651

0.642022737

0.3020314947

0.6810550117

0.6723238942

0.3490161643

0.7093239987

0.6998414752

0.3912150883

0.7317731112

0.7212685051

0.4104470826

0.7567692747

0.7477500131

0.4259269759

0.7782989259

0.7699802578

0.463205141

0.7967304158

0.7900215941

0.5034549643

0.8113869533

0.8038230464

0.5027579416

0.8244204437

0.8164248386

0.5075551632

0.8333827213

0.824222196

0.5287536667

0.8381272357

0.8318985948

0.5391145721

0.8445578678

0.8366153748

0.5497952016

0.8431383188

0.8353679574

0.546087862

0.8424371029

0.8341388571

0.5648601588

0.8292506762

0.820302969

0.5491385395

0.8245043842

0.8265647668

0.5537025935

0.8267250219

0.8172264548

0.5760861911

0.816796474

0.8095193501

0.5124813442

0.8176791988

0.8116431832

0.5302578527

0.8228844972

0.8222656797

0.5516065657

0.8348534905

0.8312283277

0.5632398511

0.8517126267

0.8502950027

0.593183815

0.8743165188

0.8750236122

0.6321456976

0.8999075075

0.8961168075

0.6439998479

0.9344449739

0.9360643344

0.7087450519

0.966487662

0.9705517464

0.749326124

1.0019690793

1.0058231237

0.7762005916

1.0389074409

1.0428406719

0.840402435

1.0664917375

1.0684646803

0.8952541385

1.0924342227

1.099481131

0.9006968864

1.1074728708

1.1092304751

0.9329864332

1.1075256705

1.1089603169

0.9014496359

1.0972329202

1.1009748612

0.9280391967

1.0716001257

1.0726912988

0.8861208759

1.0299196855

1.031996374

0.8307378714

0.9721896907

0.9714166394

0.7753728589

0.9037880604

0.9032407007

0.6900283847

0.8218846507

0.8218458554

0.6072358185

0.7372488626

0.7356345855

0.5340596263

0.6466793042

0.6477649996

0.408845934

0.5619239585

0.559812621

0.3186939535

0.4794355

0.4855430343

0.2420978086

0.4040525223

0.4021505093

0.1522980716

0.3348858815

0.3395275361

0.0824432586

0.2739228279

0.274464296

0.0069044284

0.2242297406

0.2298104462

-0.0387415345

0.1797295747

0.1830061871

-0.1162764583

0.1439705607

0.1502890248

-0.1464967464

0.1167444768

0.1233321422

-0.1643412966

0.0909937624

0.0948322817

-0.1992773314

0.0718778612

0.0771704165

-0.2144000335

0.0581230298

0.0724603555

-0.1949387131

0.0483158376

0.056375929

-0.2073451957

0.0391815479

0.0487735015

-0.2276618931

0.0302151124

0.0387286591

-0.2494427077

0.0251299025

0.0324303253

-0.2358430103

0.0210115183

0.0322025915

-0.2251604601

0.0181326329

0.0277004028

-0.2352729048

0.0159042575

0.0253564894

-0.2479640997

0.0130936941

0.022907443

-0.2554803368

0.0097902806

0.0206981633

-0.242565408

0.0104480295

0.0219716718

-0.2376716128

0.0076008913

0.0183034724

-0.2600013071

0.0075243424

0.0193988778

-0.2356601476

0.0074198354

0.0181522348

-0.2573941341

0.006159098

0.018270071

-0.2426729471

0.0044583776

0.0163686805

-0.2525288345

0.0062654036

0.0187112742

-0.2435709704

0.007223892

0.0173376396

-0.2310453268

0.00549819

0.0169670181

-0.2402852774

0.0057098512

0.0202953612

-0.2380426871

0.0049846614

0.0208967856

-0.2436032345

0.0044034177

0.0170319569

-0.2379243771

0.0051390332

0.0197951379

-0.2228794542

0.0071173307

0.0200392133

-0.2220240528

0.0046715849

0.0178033782

-0.221986391

0.0038584381

0.0142859886

-0.2584746554

0.0063180994

0.0192105005

-0.2347942427

0.0066553694

0.0213678433

-0.2162886661

0.0067085287

0.0203162425

-0.2368003344

0.0064988744

0.019832723

-0.2390752619

0.0089505999

0.0231474387

-0.1990781786

0.0088384879

0.0223225725

-0.2357247003

0.0097373086

0.0254372859

-0.2242997096

0.008657374

0.0215219284

-0.2334388561

0.0087442616

0.0210736024

-0.2283880678

0.0094146716

0.0185892563

-0.2694610771

0.0105824372

0.0170212883

-0.2596680161

0.0116415759

0.0205408457

-0.2442646621

0.0134019491

0.0222650157

-0.2610387446

0.0139620595

0.0298824069

-0.2174616278

0.0122530598

0.0233414838

-0.2228633101

0.0149621982

0.0280264798

-0.2136090386

0.0150471623

0.0289155968

-0.219985016

0.0131285808

0.0281305305

-0.2212977566

0.0152003321

0.0268885208

-0.2081847204

0.014832001

0.0322560374

-0.1872186186

0.0163060002

0.0381444568

-0.164077348

0.0150611709

0.0359568723

-0.1776716478

0.0181130982

0.0397162297

-0.1576450239

0.0168433305

0.0358834088

-0.1569554169

0.0173186981

0.039128697

-0.1315723949

0.0182902929

0.0371133709

-0.1392965863

0.0177660886

0.0345522914

-0.1442710663

0.0172582681

0.0321784235

-0.1548595592

0.0167335155

0.0299167886

-0.1879939067

0.0169584339

0.0325358288

-0.1840634412

0.0171787695

0.0298968106

-0.1811449689

0.0187058477

0.0343877199

-0.1817103698

0.0158530373

0.0312536155

-0.1776824187

0.0187104439

0.0331465699

-0.1715428669

0.0190653558

0.0352273541

-0.1538412322

0.0198483931

0.03555514

-0.1616693765

0.0182673101

0.0349177103

-0.1655802454

0.021229097

0.0333892049

-0.1823242276

0.0212700386

0.0396706413

-0.1656879789

0.0234902565

0.0375500636

-0.1793894896

0.0251727265

0.040531313

-0.1351516746

0.0253903114

0.0443864473

-0.148167274

0.0270945771

0.0410821951

-0.1540190377

0.028306217

0.0411808402

-0.1679610829

0.0304523513

0.0453528576

-0.150107172

0.0319508147

0.0480641986

-0.1317502975

0.0330426316

0.0474909323

-0.1318904534

0.0346766713

0.0518643654

-0.1348498212

0.034666977

0.0482528131

-0.1368765187

0.0355388513

0.047202686

-0.1398247873

0.0356037282

0.0502230505

-0.1221485118

0.0371979722

0.0520358328

-0.1098319531

0.0381773891

0.0498559789

-0.1382886898

0.0381457386

0.0487832844

-0.1205848547

0.0390565342

0.0494036877

-0.1336585514

0.038573613

0.0503106263

-0.1280090161

0.0374656725

0.0442933514

-0.1355505516

0.0381626049

0.0532483147

-0.1096431943

0.037310687

0.0506781849

-0.1473050768

0.0392929336

0.0496603414

-0.1250977286

0.0408235187

0.0498340859

-0.1457746201

0.0407909218

0.0494907889

-0.1758244819

0.0410149441

0.0500744501

-0.1492180596

0.0385622933

0.049583481

-0.137959909

0.0401180152

0.0520185899

-0.1504843612

0.041785592

0.0540522371

-0.1385042908

0.0408921812

0.0561321055

-0.1301923664

0.040506812

0.055072058

-0.1327691108

0.0428818302

0.0598755896

-0.112194142

0.040927321

0.0592748687

-0.1158396521

0.0406952173

0.0557750125

-0.1238954205

0.0420802067

0.0626311993

-0.1151332323

0.0411118172

0.0619406879

-0.0978087685

0.0422816351

0.0609642823

-0.1045786081

0.0441926392

0.0637179437

-0.0931151351

0.0452597081

0.0616131851

-0.0963144114

0.0449104663

0.060119996

-0.0942912879

0.0453446526

0.0618641772

-0.1323217039

0.0460462553

0.0618856371

-0.100435909

0.0480615103

0.0598965792

-0.1116278781

0.0507325463

0.0611714012

-0.1186221262

0.0540665353

0.0654503313

-0.1154567856

0.0555496528

0.0674776966

-0.1146640647

0.0576409562

0.0704209551

-0.1173926841

0.0618490661

0.0719446169

-0.1132403689

0.0634054217

0.079473914

-0.0919875123

0.0683176361

0.0811172188

-0.0905091535

0.0708359868

0.085095035

-0.1018276384

0.0748936951

0.0860805699

-0.0826633211

0.0789092275

0.0928080677

-0.0656341705

0.0800296685

0.0961203543

-0.0789449796

0.0886510572

0.094805106

-0.0803157871

0.0898312441

0.1018191891

-0.0891170777

0.091555242

0.100887456

-0.07924721

0.0960819226

0.1041886442

-0.070222684

0.1002300562

0.1117112557

-0.0634693108

0.0999182666

0.1084415976

-0.0601975252

0.1032492313

0.110944067

-0.0640847673

0.1058190614

0.1139912075

-0.0672968872

0.107042881

0.1123888328

-0.0670377649

0.1086762725

0.1161097455

-0.0777144616

0.1106630042

0.1193878044

-0.0663143803

0.1133556097

0.1184060244

-0.0589557014

0.1142929122

0.1211796679

-0.0449206931

0.1145144329

0.1212276167

-0.0652292759

0.1160307686

0.1183510486

-0.0561587668

0.1147001389

0.1159457838

-0.0575572554

0.1138452049

0.1190235636

-0.056477342

0.1134305307

0.1163488815

-0.0636258722

0.111279646

0.111213093

-0.0703576306

0.1110764441

0.1133984961

-0.0529188564

0.1070146539

0.1067948542

-0.0804614986

0.1055639481

0.1045606448

-0.0764569251

0.1056831601

0.1108097955

-0.0544416423

0.104017435

0.1043260879

-0.0627134735

0.1043833827

0.1082059917

-0.0545010404

0.1058310518

0.1028021513

-0.0445264175

0.1061047554

0.105505949

-0.0364133298

0.1077297446

0.1087283744

-0.0606942463

0.1114451815

0.112717547

-0.0325885494

0.1143822266

0.1098276908

-0.0502457571

0.1209327126

0.1179713383

-0.0418096145

0.1255847279

0.1229315089

-0.0507425876

0.1295630713

0.1283225608

-0.028255569

0.1363637494

0.1291578218

-0.0395139749

0.1425084324

0.1386546087

-0.0198531584

0.1503265352

0.1485739259

0.0135773638

0.1563282807

0.1536920136

0.0056040228

0.1608361652

0.1585229415

0.0095053504

0.1679934094

0.1665532903

0.0140965324

0.174983974

0.1736363034

0.0257360983

0.1790409664

0.1738208882

0.0121199413

0.1864317955

0.1850957511

0.0391534283

0.1883794454

0.1829493214

0.0427113091

0.1922139955

0.1862225428

0.006944982

0.1965449302

0.1913977629

0.0312134597

0.1979003881

0.1983876092

0.0406225419

0.2003898165

0.1938872356

0.0449084056

0.2024991259

0.1941496675

0.0458608762

0.2026992664

0.1961658302

0.0500200478

0.2027233787

0.1962987281

0.042362272

0.2007358982

0.1983994819

0.0459366404

0.2007380707

0.1934848457

0.0547777172

0.2008286936

0.1964985826

0.0470298443

0.1965615372

0.1921254951

0.0482042525

0.1914089239

0.1836160748

0.0219226206

0.1883799228

0.1816775592

0.0379657266

0.1853336233

0.177191123

0.0324523811

0.1791912364

0.1728303078

0.0174820828

0.1747972311

0.1671533868

0.0226068589

0.1694112466

0.1636872878

0.0208814084

0.1623625656

0.1575158925

-0.0003892552

0.1592281286

0.1500932

-0.004497856

0.1554844976

0.1540437156

0.0094620904

0.1482875381

0.1430834713

0.0132691113

0.1448624038

0.1407049493

-0.0343605225

0.1427420339

0.1341585728

-0.020226032

0.141065594

0.1349991707

-0.0331233881

0.1407231991

0.1328509927

-0.0058817219

0.1403770186

0.1346497334

-0.0147731346

0.141886993

0.1373561704

-0.0066385058

0.144410162

0.1404466804

0.0060176612

0.1497081077

0.1424197152

0.0068206171

0.1539729762

0.1497661925

-0.0094979538

0.1659884659

0.1578035518

0.0303992515

0.1764719427

0.1711878273

0.0584396754

0.1858480963

0.1796878893

0.0472273802

0.2022102554

0.1989944497

0.0514651577

0.2223498521

0.2158284555

0.0724108908

0.2455197994

0.2416434597

0.1365479689

0.2707045492

0.2649039659

0.1416753693

0.301976631

0.3010680112

0.1858313568

0.3409148172

0.3363287313

0.2316143625

0.3839418274

0.3816147487

0.2778869594

0.4303202361

0.4313994747

0.3171531999

0.4856212892

0.4858228838

0.405681955

0.5433398944

0.5433014597

0.4287048547

0.6077776567

0.6091822517

0.560672674

0.6742670894

0.6749633328

0.6268063312

0.7382618104

0.7428306809

0.6931573521

0.8042067819

0.8055762112

0.7645537392

0.8623330417

0.868318944

0.8493253362

0.9162971037

0.9124218199

0.8979285386

0.955500015

0.9519888874

0.9479721766

0.985071364

0.9827769422

0.9605304491

1

1

1

0.9984227088

0.9939251017

0.9864253213

0.9859392758

0.9805432739

0.9804142184

0.9548674664

0.9506208446

0.9498243876

0.914462658

0.9081638961

0.8819146211

0.8639602365

0.8572978709

0.8477252087

0.8018303112

0.7971200434

0.7571262225

0.731645785

0.7225972497

0.6736898804

0.6575277741

0.6492442761

0.6329187247

0.5881326835

0.5787979687

0.5084982396

0.5176099264

0.5144298567

0.4795690285

0.451038298

0.4467715218

0.4068575181

0.3904411262

0.3833261841

0.3437127576

0.3327684084

0.326454965

0.260932251

0.2837689824

0.2739100251

0.1817329936

0.239999609

0.2359327189

0.1332267173

0.2032025638

0.1980087316

0.10136864

0.1740513457

0.1691365292

0.0870634953

0.1463228203

0.1396917225

0.0478876468

0.1209239812

0.1163399562

0.0267314398

0.1048899307

0.1007299136

0.0337941338

0.0882809865

0.0864189339

-0.0101573845

0.0781418303

0.0717512232

-0.0237546635

0.0674725034

0.0568887702

-0.0374234936

0.0568176686

0.0503781742

-0.0531834554

0.0503188389

0.0481433694

-0.048047758

0.0473924222

0.0457900165

-0.0547494344

0.0428947834

0.0352571066

-0.0856151091

0.0379803306

0.0343686598

-0.0580540068

0.0371725679

0.0299915925

-0.0304761372

0.0370076659

0.0332655613

-0.0377368018

0.0323304829

0.0290902739

-0.0927482609

0.0277631323

0.0195482858

-0.0915612765

0.0278297169

0.0237885463

-0.0683495451

0.0323678508

0.0234858588

-0.0413721002

0.0269068405

0.0201259877

-0.0996267459

0.0261748711

0.0227333696

-0.0631075911

0.0283460806

0.0189144862

-0.0565097399

0.0287555826

0.0224219045

-0.0171564932

0.0249758758

0.0201060326

-0.0360351894

0.0260659398

0.0151549841

-0.0608832177

0.0258879242

0.022904658

-0.039843473

0.0257928184

0.0161947538

-0.0877789172

0.0288848695

0.0166019801

-0.0572008847

0.0262964717

0.0188003529

-0.0541014523

0.0272756417

0.0153924216

-0.0363485067

0.0245584899

0.0162856589

-0.0621573847

0.0263476013

0.0204856222

-0.0491602729

0.0268294445

0.0265522758

-0.0216202262

0.0233611507

0.0149917509

-0.0438566839

0.0255086634

0.0142243188

-0.022236251

0.0307050513

0.0139623427

-0.0899749896

0.0223664126

0.0180251235

-0.0567905196

0.0271883333

0.0156859819

-0.0398542773

0.0236909463

0.014200207

-0.03339891

0.0268043384

0.0184714159

-0.045471591

0.0244199099

0.0215817998

0.0056526861

0.0270740755

0.0191163139

0.0036521332

0.0248315228

0.0096143702

-0.038881979

0.0226494134

0.0197849283

0.0227312671

0.0252457136

0.0229914636

0.0064826708

0.0240872895

0.0212313966

-0.0113735233

0.0259980015

0.0221564024

0.0042387725

0.0263706315

0.0193603685

0.0285923581

0.0190655862

0.0063610095

-0.0237384539

0.0253852462

0.020513466

0.0145129516

0.0212964901

0.0077452123

-0.0637338489

0.0216212761

0.0077308501

-0.0246354286

0.0264595343

0.013903921

-0.0097844289

0.026182471

0.0170342767

0.0060555106

0.0173338629

0.0098724924

-0.05199004

0.0183925686

0.0106381033

-0.0739903153

0.0224973268

0.0105259453

0.0068422456

0.022346166

0.0131760046

-0.008119603

0.020363224

0.0015154643

-0.0348737311

0.0170592952

0.0028400021

-0.1170853235

0.021461641

0.0161557949

-0.0392817039

0.0260880483

0.0157578665

-0.0119086142

0.0174434649

0.0123220644

-0.0402323874

0.0166797576

0.0031492939

-0.0425874039

0.018407968

0.0026992504

-0.0681875982

0.0196187068

0.0109115545

-0.0851024748

0.0190163926

0.0082140813

-0.0574708631

0.0159127554

0.012442593

-0.0413396932

0.0162782052

0.0042152418

-0.0540852533

0.0147610476

-0.0012088267

-0.0900559235

0.0159922284

0.0022955299

-0.1022915361

0.0172879075

0.0107859512

-0.0736718616

0.0140341455

0.0023746984

-0.0850647036

0.0176691147

0.0060840131

0.0088807915

0.0132568853

0.0015460174

-0.1160769287

0.0157584088

-0.0095738211

-0.0924137482

0.0109508197

-0.0072131466

-0.0405456774

0.014745893

-0.0028468291

-0.0723926192

0.0082222757

-0.008461536

-0.1530222403

0.0100556149

0.0079838129

0.0176227021

0.0104457355

0.0046524059

-0.0474320826

0.007604558

-0.0131869521

-0.0773096764

0.0091835461

-0.0087056421

-0.0405672831

0.0085205119

-0.0182225677

-0.1279335467

0.0080109339

0.0022043253

-0.0809525971

0.0160889297

0.0010706074

-0.005957401

0.0108141057

-0.0070392808

-0.0590420877

0.0115105732

-0.0132608797

-0.0932284365

0.0106246409

0.0015946246

-0.0511206095

0.0092410968

-0.0172482165

-0.1249898977

0.0090730327

-0.010171042

-0.0481233672

0.0084063497

-0.0082932465

-0.0702010935

0.0020028828

-0.0211384236

-0.0802941967

0.0127060566

-0.0027903837

-0.0470648346

0.0003861268

-0.0125834941

-0.0746704035

0.0011003807

-0.0128228497

-0.0570227024

0.0003806366

-0.004373469

-0.1129707309

0.0120137327

-0.002014915

-0.0647380072

0.0087504512

0.0020904368

-0.0364457429

0.0130372328

0.0024043286

-0.0273154406

0.0058855574

-0.0115002944

-0.0722144976

0.0033313929

0.0076090045

0.0194021151

0.009420175

-0.0005026208

-0.0085087887

0.0099159506

0.0198508193

0.0042360691

0.0006572165

-0.0030531742

-0.0610829844

0.0014694565

-0.00608295

-0.0855989223

0.0114213288

-0.0038257637

-0.0615041084

0.0110756102

0.0038202405

-0.018707501

0.0105234895

0.0087436773

0.0159488193

0.0068620534

-0.0004628196

-0.0691916592

0.0080943693

-0.007453593

-0.06318317

0.0108136469

0.0032377314

0.0271398629

0.0082218174

-0.0035380177

-0.0829115631

-0.0033248468

-0.0258745863

-0.0972908764

0.0068826764

-0.0107451065

0.0686016104

0.0131680586

-0.0052587832

0.0250085362

0.0019738181

0.0140736234

0.0841737617

0.0080893262

-0.0051847796

0.0037116078

0.0072051005

0.0225398052

0.1262921955

0.0055627867

0.0010289468

-0.0273694715

0.0048744977

-0.0055409151

-0.0477993306

0.0071113726

-0.0040496615

-0.0188858368

0.0013248423

-0.0136702395

-0.082031927

0.0045683006

0.0142340559

-0.1008944273

0.0010475278

-0.0055612651

-0.0341120161

0.0101714311

-0.0161545825

-0.0645112585

0.0051136098

0.0068270155

-0.0456498188

-0.0008713875

0.0056523264

-0.07326704

0.0118948767

-0.0134512396

0.0540036812

0.0075855356

0.0184032186

0.1060883498

0.0010617134

-0.0057074146

0.067838169

0.0045857054

-0.0091744149

0.0550970765

-0.0039258559

-0.0241140884

-0.0423119368

-0.000032023

-0.0162811338

-0.0042167531

0.0039385771

-0.0206731419

-0.0643924926

-0.0107241845

-0.0123228751

-0.1046972707

0.0026712003

0.0093038938

0.0400002593

0.0036967909

-0.0036916069

-0.0012975043

0.004703418

0.0021552512

0.0289494035

-0.0031717231

-0.0056602401

-0.0835753186

-0.0054222568

-0.0032817207

-0.0754908055

0.0019639773

-0.0279704216

-0.0634099187

-0.0024489874

-0.043202648

-0.0505859805

-0.0131353382

-0.0192309593

-0.1956708204

-0.0045153476

-0.0375780094

0.0134502781

0.0107838275

-0.0023351059

-0.0185453779

-0.0074162468

-0.038204399

-0.1817642244

-0.0017160557

-0.0025053757

0.0384662345

0.0045708193

-0.0184931465

-0.0818322514

-0.0048292496

-0.0256374776

-0.0899048448

-0.0164770711

-0.0413475577

-0.2420545303

-0.0170373112

-0.0372141266

-0.0981540279

-0.0123109644

-0.0421126607

-0.258356394

-0.0113103525

-0.0289747835

-0.0949117226

-0.0062060923

-0.037426932

-0.1289039273

-0.021195564

-0.0471332398

-0.3289694695

-0.0151298817

-0.0337835875

-0.1762014644

0.0053442616

-0.0356718148

-0.0406969203

-0.0122416507

-0.0703393697

-0.2216044241

-0.0232864148

-0.0616153498

-0.2828733895

-0.0007561287

-0.0125049393

0.0636990197

-0.0032041767

-0.0315925206

-0.1768692378

-0.005567982

-0.0688531702

-0.4098518367

-0.0116054548

-0.0139132585

-0.1976678918

-0.0060681563

-0.0317354157

-0.1385635796

-0.0000699928

-0.0383978301

-0.2295930018

-0.0101789909

-0.0646554133

-0.2226265928

-0.0066267217

-0.0286307871

-0.2130171238

-0.014418055

-0.0227731506

-0.0661200355

-0.007829895

-0.047554617

-0.1179588873

0.008917356

-0.0207589985

-0.0324372793

-0.0068459236

-0.0443900399

-0.2002838784



								In vitro		In vivo				118.2321279285		0.054543378

		Blank		Scale		Fo		Chl		Chl				chl		Fo

		0		1		5		591.22		10304.43				127.59		1.0759493671

		0		3.16		5		187.13		3260.43				63.79		0.545

		0		10		5		59.17		1029.83				12.76		0.1060126582

		0.05		31.6		5		18.58		322.17				6.38		0.05

		0.1		100		5		5.85		100.31				1.28		0.0107594937

		0.4		316		5		1.78		29.32				0.64		0.0054

		1.3		1000		5		0.49		6.95				0.13		0.0010126582

		4.2		3160		5		0.08		-0.16				0.06		0.0005063291





		



Chl (ug/L)

Fo



		





		






_1211715664.xls
Chart2

		2

		4

		10

		20

		1.1

		2.2

		5.5

		11

		1.1

		2.2

		5.5

		11



Fo

Chl+Pheo (mg/m3)

Fa

0.4

0.8

2

4

0.22

0.44

1.1

2.2

0.22

0.44

1.1

2.2



Sheet1

		1		0.2				0.8		0		0.2

		Chl		Pheo		Chl+Pheo		Chl/Pheo		Fo		Fa		Fo-Fa		Fo/Fa

		1		1		2		1		1.2		0.4		0.8		3.0

		2		2		4		1		2.4		0.8		1.6		3.0

		5		5		10		1		6.0		2.0		4.0		3.0

		10		10		20		1		12.0		4.0		8.0		3.0

		0.1		1		1.1		0.1		0.3		0.2		0.1		1.4

		0.2		2		2.2		0.1		0.6		0.4		0.2		1.4

		0.5		5		5.5		0.1		1.5		1.1		0.4		1.4

		1		10		11		0.1		3.0		2.2		0.8		1.4

		1		0.1		1.1		10		1.0		0.2		0.8		4.6

		2		0.2		2.2		10		2.0		0.4		1.6		4.6

		5		0.5		5.5		10		5.1		1.1		4.0		4.6

		10		1		11		10		10.2		2.2		8.0		4.6

												5





Sheet1

		



Fo

Chl (mg/m3)

Fo-Fa



Sheet2

		



Pheo (mg/m3)

Fo/Fa



Sheet3

		



Fo

Chl+Pheo (mg/m3)

Fa



Sheet4

		



Chl/Pheo (mg/m3)

Fo/Fa



Sheet5

		





		





		





		






_1211717871.xls
Chart3

		1

		2

		5

		10

		1

		2

		5

		10

		0.1

		0.2

		0.5

		1



Pheo (mg/m3)

A*Fa-Fo

0.8

1.6

4

8

0.8

1.6

4

8

0.08

0.16

0.4

0.8



Sheet1

		1		0.2				0.8		0		0.2				0.8		0

		Chl		Pheo		Chl+Pheo		Chl/Pheo		Fo		Fa		Fo-Fa		Fo/Fa

		1		1		2		1		1.2		0.4		0.8		3.0		0.8

		2		2		4		1		2.4		0.8		1.6		3.0		1.6

		5		5		10		1		6.0		2.0		4.0		3.0		4

		10		10		20		1		12.0		4.0		8.0		3.0		8

		0.1		1		1.1		0.1		0.3		0.2		0.1		1.4		0.8

		0.2		2		2.2		0.1		0.6		0.4		0.2		1.4		1.6

		0.5		5		5.5		0.1		1.5		1.1		0.4		1.4		4

		1		10		11		0.1		3.0		2.2		0.8		1.4		8

		1		0.1		1.1		10		1.0		0.2		0.8		4.6		0.08

		2		0.2		2.2		10		2.0		0.4		1.6		4.6		0.16

		5		0.5		5.5		10		5.1		1.1		4.0		4.6		0.4

		10		1		11		10		10.2		2.2		8.0		4.6		0.8

												5





Sheet1

		



Fo

Chl (mg/m3)

Fo-Fa



Sheet2

		



Pheo (mg/m3)

A*Fa-Fo



Sheet3

		



Fo

Chl+Pheo (mg/m3)

Fa



Sheet4

		



Chl/Pheo (mg/m3)

Fo/Fa



Sheet5

		





		





		





		






_1211715657.xls
Chart1

		1

		2

		5

		10

		0.1

		0.2

		0.5

		1

		1

		2

		5

		10



Fo

Chl (mg/m3)

Fo-Fa

0.8

1.6

4

8

0.08

0.16

0.4

0.8

0.8

1.6

4

8



Sheet1

		1		0.2				0.8		0		0.2

		Chl		Pheo		Chl+Pheo		Chl/Pheo		Fo		Fa		Fo-Fa		Fo/Fa

		1		1		2		1		1.2		0.4		0.8		3.0

		2		2		4		1		2.4		0.8		1.6		3.0

		5		5		10		1		6.0		2.0		4.0		3.0

		10		10		20		1		12.0		4.0		8.0		3.0

		0.1		1		1.1		0.1		0.3		0.2		0.1		1.4

		0.2		2		2.2		0.1		0.6		0.4		0.2		1.4

		0.5		5		5.5		0.1		1.5		1.1		0.4		1.4

		1		10		11		0.1		3.0		2.2		0.8		1.4

		1		0.1		1.1		10		1.0		0.2		0.8		4.6

		2		0.2		2.2		10		2.0		0.4		1.6		4.6

		5		0.5		5.5		10		5.1		1.1		4.0		4.6

		10		1		11		10		10.2		2.2		8.0		4.6

												5





Sheet1

		



Fo

Chl (mg/m3)

Fo-Fa



Sheet2

		



Pheo (mg/m3)

Fo/Fa



Sheet3

		



Fo

Chl+Pheo (mg/m3)

Fa



Sheet4

		



Chl/Pheo (mg/m3)

Fo/Fa



Sheet5

		





		





		





		






